## README
sheet | description | sign_convention | unit
README | Defines output tables, sign conventions and units. | Differences are automatic minus manual unless explicitly labelled otherwise. | See table column name; indentation dimensions are in um.
Data_Quality | Counts, missing values, image matches and calibration assets found in the source data. | Not applicable. | Counts or text.
Bland_Altman | Method-specific Bland-Altman bias, limits of agreement and bootstrap confidence intervals. | Automatic minus manual. | um, percent, or log ratio depending on representation.
Deming | Error-in-variables regression results and lambda sensitivity. | x = manual, y = automatic. | Same unit as fitted measurand.
En_Results | Normalized error from series mean hardness and expanded uncertainties. | Automatic minus manual in numerator. | Dimensionless.
Image_Measurements | Independent Python image-processing measurements where image and calibration were available. | Python minus comparator where labelled. | um and hardness units.

## Data_Quality
category | method | item | value
workbook | Vickers | valid_observations | 60
workbook | Vickers | series_count | 20
workbook | Vickers | complete_three_position_series | 20
workbook | Vickers | incomplete_series | 0
workbook | Vickers | magnification_20X | 54
workbook | Vickers | magnification_50X | 6
images | Vickers | picture_ids_in_workbook | 8
missing_values | Vickers | hardness_scale_raw | 0
missing_values | Vickers | applied_force_n | 0
missing_values | Vickers | manual_d1_um | 0
missing_values | Vickers | manual_d2_um | 0
missing_values | Vickers | manual_mean_dimension_um | 0
missing_values | Vickers | manual_series_mean_dimension_um | 40
missing_values | Vickers | manual_hardness | 0
missing_values | Vickers | manual_series_mean_hardness | 40
missing_values | Vickers | manual_expanded_uncertainty | 40
missing_values | Vickers | manual_relative_expanded_uncertainty | 40
missing_values | Vickers | automatic_d1_um | 0
missing_values | Vickers | automatic_d2_um | 0
missing_values | Vickers | automatic_mean_dimension_um | 0
missing_values | Vickers | automatic_series_mean_dimension_um | 40
missing_values | Vickers | automatic_hardness | 0
missing_values | Vickers | automatic_series_mean_hardness | 40
missing_values | Vickers | automatic_expanded_uncertainty | 40
missing_values | Vickers | automatic_relative_expanded_uncertainty | 40
missing_values | Vickers | hardness_scale | 0
missing_values | Vickers | hardness_level | 0
missing_values | Vickers | ball_diameter_mm | 60
workbook | Micro-Vickers | valid_observations | 54
workbook | Micro-Vickers | series_count | 18
workbook | Micro-Vickers | complete_three_position_series | 18
workbook | Micro-Vickers | incomplete_series | 0
workbook | Micro-Vickers | magnification_50X | 45
workbook | Micro-Vickers | magnification_20X | 9
images | Micro-Vickers | picture_ids_in_workbook | 20
missing_values | Micro-Vickers | hardness_scale_raw | 0
missing_values | Micro-Vickers | applied_force_n | 0
missing_values | Micro-Vickers | manual_d1_um | 0
missing_values | Micro-Vickers | manual_d2_um | 0
missing_values | Micro-Vickers | manual_mean_dimension_um | 0
missing_values | Micro-Vickers | manual_series_mean_dimension_um | 36
missing_values | Micro-Vickers | manual_hardness | 0
missing_values | Micro-Vickers | manual_series_mean_hardness | 36
missing_values | Micro-Vickers | manual_expanded_uncertainty | 36
missing_values | Micro-Vickers | manual_relative_expanded_uncertainty | 36
missing_values | Micro-Vickers | automatic_d1_um | 0
missing_values | Micro-Vickers | automatic_d2_um | 0
missing_values | Micro-Vickers | automatic_mean_dimension_um | 0
missing_values | Micro-Vickers | automatic_series_mean_dimension_um | 36
missing_values | Micro-Vickers | automatic_hardness | 0
missing_values | Micro-Vickers | automatic_series_mean_hardness | 36
missing_values | Micro-Vickers | automatic_expanded_uncertainty | 36
missing_values | Micro-Vickers | automatic_relative_expanded_uncertainty | 36
missing_values | Micro-Vickers | hardness_scale | 0
missing_values | Micro-Vickers | hardness_level | 0
missing_values | Micro-Vickers | ball_diameter_mm | 54
workbook | Brinell | valid_observations | 36
workbook | Brinell | series_count | 12
workbook | Brinell | complete_three_position_series | 12
workbook | Brinell | incomplete_series | 0
workbook | Brinell | magnification_20X | 33
workbook | Brinell | magnification_10X | 3
images | Brinell | picture_ids_in_workbook | 3
missing_values | Brinell | hardness_scale_raw | 0
missing_values | Brinell | ball_diameter_mm | 0
missing_values | Brinell | applied_force_n | 0
missing_values | Brinell | manual_d1_um | 0
missing_values | Brinell | manual_d2_um | 0
missing_values | Brinell | manual_mean_dimension_um | 0
missing_values | Brinell | manual_series_mean_dimension_um | 24
missing_values | Brinell | manual_hardness | 0
missing_values | Brinell | manual_series_mean_hardness | 24
missing_values | Brinell | manual_expanded_uncertainty | 24
missing_values | Brinell | manual_relative_expanded_uncertainty | 24
missing_values | Brinell | automatic_d1_um | 0
missing_values | Brinell | automatic_d2_um | 0
missing_values | Brinell | automatic_mean_dimension_um | 0
missing_values | Brinell | automatic_series_mean_dimension_um | 24
missing_values | Brinell | automatic_hardness | 0
missing_values | Brinell | automatic_series_mean_hardness | 24
missing_values | Brinell | automatic_expanded_uncertainty | 24
missing_values | Brinell | automatic_relative_expanded_uncertainty | 24
missing_values | Brinell | hardness_scale | 0
missing_values | Brinell | hardness_level | 0
workbook | Knoop | valid_observations | 33
workbook | Knoop | series_count | 11
workbook | Knoop | complete_three_position_series | 11
workbook | Knoop | incomplete_series | 0
workbook | Knoop | magnification_50X | 9
workbook | Knoop | magnification_10X | 6
workbook | Knoop | magnification_20X | 18
images | Knoop | picture_ids_in_workbook | 0
missing_values | Knoop | hardness_scale_raw | 0
missing_values | Knoop | applied_force_n | 0
missing_values | Knoop | manual_d_um | 0
missing_values | Knoop | manual_hardness | 33
missing_values | Knoop | automatic_d_um | 33
missing_values | Knoop | automatic_hardness | 33
missing_values | Knoop | delta_um | 33
missing_values | Knoop | hardness_scale | 0
missing_values | Knoop | hardness_level | 0
missing_values | Knoop | ball_diameter_mm | 33
images | All | duplicate_picture_ids | 0
calibration | All | calibration_image | stage_micrometer_0.002mm_20x.bmp
calibration | All | calibration_image | stage_micrometer_0.002mm_50x.bmp

## Clean_Vickers
method | original_sheet | excel_row | serial_number | hardness_scale_raw | position | applied_force_n | magnification_raw | manual_d1_um | manual_d2_um | manual_mean_dimension_um | manual_series_mean_dimension_um | manual_hardness | manual_series_mean_hardness | manual_expanded_uncertainty | manual_relative_expanded_uncertainty | automatic_d1_um | automatic_d2_um | automatic_mean_dimension_um | automatic_series_mean_dimension_um | automatic_hardness | automatic_series_mean_hardness | automatic_expanded_uncertainty | automatic_relative_expanded_uncertainty | picture_id | magnification | hardness_scale | hardness_level | ball_diameter_mm | series_id | record_id
Vickers | Vickers | 3 | 293-250 | HV 3 | 1 | 29.43 | 20X | 183.943 | 183.904 | 183.9235 | 184.5562 | 164.421 | 163.5923 | 2.5363 | 1.5504 | 184.291 | 184.428 | 184.3595 | 184.5995 | 163.709 | 163.2407 | 2.0781 | 1.273 | 293-250_160_HV_3_1_20X | 20X | HV 3 | 3 |  | Vickers¦SN=293-250¦scale=HV 3¦F=29.43N¦mag=20X | Vickers¦293-250¦HV 3¦1¦row3
Vickers | Vickers | 4 | 293-250 | HV 3 | 2 | 29.43 | 20X | 184.644 | 185.667 | 185.1555 |  | 163.12 |  |  |  | 184.721 | 184.703 | 184.712 |  | 163.02 |  |  |  |  | 20X | HV 3 | 3 |  | Vickers¦SN=293-250¦scale=HV 3¦F=29.43N¦mag=20X | Vickers¦293-250¦HV 3¦2¦row4
Vickers | Vickers | 5 | 293-250 | HV 3 | 3 | 29.43 | 20X | 184.115 | 185.064 | 184.5895 |  | 163.236 |  |  |  | 184.334 | 185.12 | 184.727 |  | 162.993 |  |  |  |  | 20X | HV 3 | 3 |  | Vickers¦SN=293-250¦scale=HV 3¦F=29.43N¦mag=20X | Vickers¦293-250¦HV 3¦3¦row5
Vickers | Vickers | 6 | 294-367 | HV 3 | 1 | 29.43 | 20X | 135.504 | 135.211 | 135.3575 | 135.062 | 303.553 | 304.9003 | 4.2319 | 1.388 | 135.698 | 135.197 | 135.4475 | 135.0908 | 303.17 | 304.4983 | 4.3009 | 1.4124 | 294-367_300_HV_3_1_20X | 20X | HV 3 | 3 |  | Vickers¦SN=294-367¦scale=HV 3¦F=29.43N¦mag=20X | Vickers¦294-367¦HV 3¦1¦row6
Vickers | Vickers | 7 | 294-367 | HV 3 | 2 | 29.43 | 20X | 135.058 | 134.521 | 134.7895 |  | 306.139 |  |  |  | 135.251 | 134.327 | 134.789 |  | 306.141 |  |  |  |  | 20X | HV 3 | 3 |  | Vickers¦SN=294-367¦scale=HV 3¦F=29.43N¦mag=20X | Vickers¦294-367¦HV 3¦2¦row7
Vickers | Vickers | 8 | 294-367 | HV 3 | 3 | 29.43 | 20X | 135.182 | 134.896 | 135.039 |  | 305.009 |  |  |  | 134.85 | 135.222 | 135.036 |  | 304.184 |  |  |  |  | 20X | HV 3 | 3 |  | Vickers¦SN=294-367¦scale=HV 3¦F=29.43N¦mag=20X | Vickers¦294-367¦HV 3¦3¦row8
Vickers | Vickers | 9 | 294-481 | HV 3 | 1 | 29.43 | 20X | 103.915 | 103.59 | 103.7525 | 103.873 | 516.694 | 515.499 | 7.4144 | 1.4383 | 103.96 | 104.292 | 104.126 | 104.2282 | 512.994 | 511.9953 | 7.4379 | 1.4527 | 294_481_503_HV_3_1_20X | 20X | HV 3 | 3 |  | Vickers¦SN=294-481¦scale=HV 3¦F=29.43N¦mag=20X | Vickers¦294-481¦HV 3¦1¦row9
Vickers | Vickers | 10 | 294-481 | HV 3 | 2 | 29.43 | 20X | 103.585 | 104.006 | 103.7955 |  | 516.266 |  |  |  | 103.962 | 104.188 | 104.075 |  | 513.497 |  |  |  |  | 20X | HV 3 | 3 |  | Vickers¦SN=294-481¦scale=HV 3¦F=29.43N¦mag=20X | Vickers¦294-481¦HV 3¦2¦row10
Vickers | Vickers | 11 | 294-481 | HV 3 | 3 | 29.43 | 20X | 103.88 | 104.262 | 104.071 |  | 513.537 |  |  |  | 104.704 | 104.263 | 104.4835 |  | 509.495 |  |  |  |  | 20X | HV 3 | 3 |  | Vickers¦SN=294-481¦scale=HV 3¦F=29.43N¦mag=20X | Vickers¦294-481¦HV 3¦3¦row11
Vickers | Vickers | 12 | 216-610 | HV 3 | 1 | 29.43 | 50X | 86.998 | 87.315 | 87.1565 | 87.4893 | 732.203 | 726.6623 | 11.4188 | 1.5714 | 86.496 | 86.765 | 86.6305 | 87.2083 | 740.873 | 731.3793 | 14.3219 | 1.9582 | 216-610_701_HV_3_1_50X | 50X | HV 3 | 3 |  | Vickers¦SN=216-610¦scale=HV 3¦F=29.43N¦mag=50X | Vickers¦216-610¦HV 3¦1¦row12
Vickers | Vickers | 13 | 216-610 | HV 3 | 2 | 29.43 | 50X | 87.835 | 87.775 | 87.805 |  | 721.427 |  |  |  | 87.981 | 88.341 | 88.161 |  | 715.613 |  |  |  |  | 50X | HV 3 | 3 |  | Vickers¦SN=216-610¦scale=HV 3¦F=29.43N¦mag=50X | Vickers¦216-610¦HV 3¦2¦row13
Vickers | Vickers | 14 | 216-610 | HV 3 | 3 | 29.43 | 50X | 87.337 | 87.676 | 87.5065 |  | 726.357 |  |  |  | 86.622 | 87.045 | 86.8335 |  | 737.652 |  |  |  |  | 50X | HV 3 | 3 |  | Vickers¦SN=216-610¦scale=HV 3¦F=29.43N¦mag=50X | Vickers¦216-610¦HV 3¦3¦row14
Vickers | Vickers | 15 | 221-101 | HV 3 | 1 | 29.43 | 50X | 78.239 | 78.245 | 78.242 | 78.1743 | 908.555 | 910.1297 | 14.1415 | 1.5538 | 78.166 | 77.845 | 78.0055 | 77.9745 | 914.084 | 914.8567 | 15.3644 | 1.6794 | 221-101_898_HV_3_1_50X | 50X | HV 3 | 3 |  | Vickers¦SN=221-101¦scale=HV 3¦F=29.43N¦mag=50X | Vickers¦221-101¦HV 3¦1¦row15
Vickers | Vickers | 16 | 221-101 | HV 3 | 2 | 29.43 | 50X | 78.122 | 78.238 | 78.18 |  | 909.996 |  |  |  | 77.253 | 77.823 | 77.538 |  | 925.128 |  |  |  |  | 50X | HV 3 | 3 |  | Vickers¦SN=221-101¦scale=HV 3¦F=29.43N¦mag=50X | Vickers¦221-101¦HV 3¦2¦row16
Vickers | Vickers | 17 | 221-101 | HV 3 | 3 | 29.43 | 50X | 78.092 | 78.11 | 78.101 |  | 911.838 |  |  |  | 78.291 | 78.469 | 78.38 |  | 905.358 |  |  |  |  | 50X | HV 3 | 3 |  | Vickers¦SN=221-101¦scale=HV 3¦F=29.43N¦mag=50X | Vickers¦221-101¦HV 3¦3¦row17
Vickers | Vickers | 18 | 293-250 | HV 20 | 1 | 196.2 | 20X | 478.702 | 478.167 | 478.4345 | 479.2792 | 161.992 | 161.4267 | 4.9143 | 3.0443 | 478.184 | 477.252 | 477.718 | 479.3273 | 162.479 | 161.3973 | 6.1848 | 3.832 |  | 20X | HV 20 | 20 |  | Vickers¦SN=293-250¦scale=HV 20¦F=196.2N¦mag=20X | Vickers¦293-250¦HV 20¦1¦row18
Vickers | Vickers | 19 | 293-250 | HV 20 | 2 | 196.2 | 20X | 481.128 | 481.706 | 481.417 |  | 159.991 |  |  |  | 481.719 | 482.478 | 482.0985 |  | 159.539 |  |  |  |  | 20X | HV 20 | 20 |  | Vickers¦SN=293-250¦scale=HV 20¦F=196.2N¦mag=20X | Vickers¦293-250¦HV 20¦2¦row19
Vickers | Vickers | 20 | 293-250 | HV 20 | 3 | 196.2 | 20X | 478.297 | 477.675 | 477.986 |  | 162.297 |  |  |  | 478.515 | 477.816 | 478.1655 |  | 162.174 |  |  |  |  | 20X | HV 20 | 20 |  | Vickers¦SN=293-250¦scale=HV 20¦F=196.2N¦mag=20X | Vickers¦293-250¦HV 20¦3¦row20
Vickers | Vickers | 21 | 294-367 | HV 20 | 1 | 196.2 | 20X | 349.164 | 349.313 | 349.2385 | 349.196 | 304.015 | 304.09 | 3.7984 | 1.2491 | 349.57 | 350.015 | 349.7925 | 349.4752 | 303.054 | 303.6047 | 3.8082 | 1.2543 |  | 20X | HV 20 | 20 |  | Vickers¦SN=294-367¦scale=HV 20¦F=196.2N¦mag=20X | Vickers¦294-367¦HV 20¦1¦row21
Vickers | Vickers | 22 | 294-367 | HV 20 | 2 | 196.2 | 20X | 349.606 | 349.457 | 349.5315 |  | 303.506 |  |  |  | 349.852 | 349.266 | 349.559 |  | 303.458 |  |  |  |  | 20X | HV 20 | 20 |  | Vickers¦SN=294-367¦scale=HV 20¦F=196.2N¦mag=20X | Vickers¦294-367¦HV 20¦2¦row22
Vickers | Vickers | 23 | 294-367 | HV 20 | 3 | 196.2 | 20X | 348.764 | 348.872 | 348.818 |  | 304.749 |  |  |  | 348.974 | 349.174 | 349.074 |  | 304.302 |  |  |  |  | 20X | HV 20 | 20 |  | Vickers¦SN=294-367¦scale=HV 20¦F=196.2N¦mag=20X | Vickers¦294-367¦HV 20¦3¦row23
Vickers | Vickers | 24 | 294-481 | HV 20 | 1 | 196.2 | 20X | 270.186 | 269.66 | 269.923 | 270.0955 | 508.932 | 508.284 | 6.497 | 1.2782 | 270.038 | 269.777 | 269.9075 | 269.9183 | 508.993 | 508.951 | 6.313 | 1.2404 | 294_481_503_HV_20_1_20X | 20X | HV 20 | 20 |  | Vickers¦SN=294-481¦scale=HV 20¦F=196.2N¦mag=20X | Vickers¦294-481¦HV 20¦1¦row24
Vickers | Vickers | 25 | 294-481 | HV 20 | 2 | 196.2 | 20X | 269.873 | 269.85 | 269.8615 |  | 509.164 |  |  |  | 269.547 | 269.722 | 269.6345 |  | 510.02 |  |  |  |  | 20X | HV 20 | 20 |  | Vickers¦SN=294-481¦scale=HV 20¦F=196.2N¦mag=20X | Vickers¦294-481¦HV 20¦2¦row25
Vickers | Vickers | 26 | 294-481 | HV 20 | 3 | 196.2 | 20X | 270.479 | 270.525 | 270.502 |  | 506.756 |  |  |  | 270.397 | 270.029 | 270.213 |  | 507.84 |  |  |  |  | 20X | HV 20 | 20 |  | Vickers¦SN=294-481¦scale=HV 20¦F=196.2N¦mag=20X | Vickers¦294-481¦HV 20¦3¦row26
Vickers | Vickers | 27 | 216-610 | HV 20 | 1 | 196.2 | 20X | 227.965 | 227.478 | 227.7215 | 227.4537 | 715.042 | 716.7313 | 9.5727 | 1.3356 | 227.615 | 227.823 | 227.719 | 227.1698 | 715.058 | 717.44 | 10.0422 | 1.3997 | 216-610_701_HV_20_1_20X | 20X | HV 20 | 20 |  | Vickers¦SN=216-610¦scale=HV 20¦F=196.2N¦mag=20X | Vickers¦216-610¦HV 20¦1¦row27
Vickers | Vickers | 28 | 216-610 | HV 20 | 2 | 196.2 | 20X | 227.757 | 227.532 | 227.6445 |  | 715.526 |  |  |  | 227.073 | 226.9 | 226.9865 |  | 716.423 |  |  |  |  | 20X | HV 20 | 20 |  | Vickers¦SN=216-610¦scale=HV 20¦F=196.2N¦mag=20X | Vickers¦216-610¦HV 20¦2¦row28
Vickers | Vickers | 29 | 216-610 | HV 20 | 3 | 196.2 | 20X | 227.274 | 226.716 | 226.995 |  | 719.626 |  |  |  | 226.263 | 227.345 | 226.804 |  | 720.839 |  |  |  |  | 20X | HV 20 | 20 |  | Vickers¦SN=216-610¦scale=HV 20¦F=196.2N¦mag=20X | Vickers¦216-610¦HV 20¦3¦row29
Vickers | Vickers | 30 | 221-101 | HV 20 | 1 | 196.2 | 20X | 202.798 | 202.576 | 202.687 | 202.8217 | 902.585 | 901.3887 | 11.2763 | 1.251 | 202.346 | 203.176 | 202.761 | 202.566 | 901.926 | 903.6693 | 12.0942 | 1.3383 | 221-101_898_HV_20_1_20X | 20X | HV 20 | 20 |  | Vickers¦SN=221-101¦scale=HV 20¦F=196.2N¦mag=20X | Vickers¦221-101¦HV 20¦1¦row30
Vickers | Vickers | 31 | 221-101 | HV 20 | 2 | 196.2 | 20X | 203.287 | 202.824 | 203.0555 |  | 899.312 |  |  |  | 203.227 | 202.373 | 202.8 |  | 901.579 |  |  |  |  | 20X | HV 20 | 20 |  | Vickers¦SN=221-101¦scale=HV 20¦F=196.2N¦mag=20X | Vickers¦221-101¦HV 20¦2¦row31
Vickers | Vickers | 32 | 221-101 | HV 20 | 3 | 196.2 | 20X | 202.939 | 202.506 | 202.7225 |  | 902.269 |  |  |  | 202.175 | 202.099 | 202.137 |  | 907.503 |  |  |  |  | 20X | HV 20 | 20 |  | Vickers¦SN=221-101¦scale=HV 20¦F=196.2N¦mag=20X | Vickers¦221-101¦HV 20¦3¦row32
Vickers | Vickers | 33 | 293-250 | HV 50 | 1 | 490.5 | 20X | 756.607 | 756.335 | 756.471 | 756.2575 | 162 | 162.1 | 6.9689 | 4.2991 | 756.809 | 757.33 | 757.0695 | 756.645 | 162 | 162.0333 | 7.3001 | 4.5053 |  | 20X | HV 50 | 50 |  | Vickers¦SN=293-250¦scale=HV 50¦F=490.5N¦mag=20X | Vickers¦293-250¦HV 50¦1¦row33
Vickers | Vickers | 34 | 293-250 | HV 50 | 2 | 490.5 | 20X | 758.557 | 759.2 | 758.8785 |  | 161 |  |  |  | 758.982 | 759.583 | 759.2825 |  | 160.8 |  |  |  |  | 20X | HV 50 | 50 |  | Vickers¦SN=293-250¦scale=HV 50¦F=490.5N¦mag=20X | Vickers¦293-250¦HV 50¦2¦row34
Vickers | Vickers | 35 | 293-250 | HV 50 | 3 | 490.5 | 20X | 753.114 | 753.732 | 753.423 |  | 163.3 |  |  |  | 753.544 | 753.622 | 753.583 |  | 163.3 |  |  |  |  | 20X | HV 50 | 50 |  | Vickers¦SN=293-250¦scale=HV 50¦F=490.5N¦mag=20X | Vickers¦293-250¦HV 50¦3¦row35
Vickers | Vickers | 36 | 294-367 | HV 50 | 1 | 490.5 | 20X | 550.791 | 552.234 | 551.5125 | 551.6933 | 304.8 | 304.6333 | 3.8664 | 1.2692 | 550.99 | 552.97 | 551.98 | 551.7327 | 304.3 | 304.5667 | 3.851 | 1.2644 |  | 20X | HV 50 | 50 |  | Vickers¦SN=294-367¦scale=HV 50¦F=490.5N¦mag=20X | Vickers¦294-367¦HV 50¦1¦row36
Vickers | Vickers | 37 | 294-367 | HV 50 | 2 | 490.5 | 20X | 552.214 | 552.175 | 552.1945 |  | 304.1 |  |  |  | 552.339 | 551.633 | 551.986 |  | 304.3 |  |  |  |  | 20X | HV 50 | 50 |  | Vickers¦SN=294-367¦scale=HV 50¦F=490.5N¦mag=20X | Vickers¦294-367¦HV 50¦2¦row37
Vickers | Vickers | 38 | 294-367 | HV 50 | 3 | 490.5 | 20X | 550.944 | 551.802 | 551.373 |  | 305 |  |  |  | 550.974 | 551.49 | 551.232 |  | 305.1 |  |  |  |  | 20X | HV 50 | 50 |  | Vickers¦SN=294-367¦scale=HV 50¦F=490.5N¦mag=20X | Vickers¦294-367¦HV 50¦3¦row38
Vickers | Vickers | 39 | 294-481 | HV 50 | 1 | 490.5 | 20X | 426.107 | 426.543 | 426.325 | 426.5372 | 510.1 | 509.6 | 5.8116 | 1.1404 | 426.539 | 426.462 | 426.5005 | 426.8152 | 509.7 | 508.9333 | 6.2624 | 1.2305 |  | 20X | HV 50 | 50 |  | Vickers¦SN=294-481¦scale=HV 50¦F=490.5N¦mag=20X | Vickers¦294-481¦HV 50¦1¦row39
Vickers | Vickers | 40 | 294-481 | HV 50 | 2 | 490.5 | 20X | 426.331 | 427.089 | 426.71 |  | 509.2 |  |  |  | 426.734 | 427.684 | 427.209 |  | 508 |  |  |  |  | 20X | HV 50 | 50 |  | Vickers¦SN=294-481¦scale=HV 50¦F=490.5N¦mag=20X | Vickers¦294-481¦HV 50¦2¦row40
Vickers | Vickers | 41 | 294-481 | HV 50 | 3 | 490.5 | 20X | 426.65 | 426.503 | 426.5765 |  | 509.5 |  |  |  | 427.247 | 426.225 | 426.736 |  | 509.1 |  |  |  |  | 20X | HV 50 | 50 |  | Vickers¦SN=294-481¦scale=HV 50¦F=490.5N¦mag=20X | Vickers¦294-481¦HV 50¦3¦row41
Vickers | Vickers | 42 | 216-610 | HV 50 | 1 | 490.5 | 20X | 360.468 | 360.302 | 360.385 | 360.7707 | 713.9 | 712.3667 | 9.0466 | 1.2699 | 360.391 | 360.031 | 360.211 | 360.494 | 714.6 | 713.4667 | 8.4852 | 1.1893 |  | 20X | HV 50 | 50 |  | Vickers¦SN=216-610¦scale=HV 50¦F=490.5N¦mag=20X | Vickers¦216-610¦HV 50¦1¦row42
Vickers | Vickers | 43 | 216-610 | HV 50 | 2 | 490.5 | 20X | 360.678 | 360.832 | 360.755 |  | 712.4 |  |  |  | 360.795 | 360.281 | 360.538 |  | 713.3 |  |  |  |  | 20X | HV 50 | 50 |  | Vickers¦SN=216-610¦scale=HV 50¦F=490.5N¦mag=20X | Vickers¦216-610¦HV 50¦2¦row43
Vickers | Vickers | 44 | 216-610 | HV 50 | 3 | 490.5 | 20X | 361.138 | 361.206 | 361.172 |  | 710.8 |  |  |  | 361.025 | 360.441 | 360.733 |  | 712.5 |  |  |  |  | 20X | HV 50 | 50 |  | Vickers¦SN=216-610¦scale=HV 50¦F=490.5N¦mag=20X | Vickers¦216-610¦HV 50¦3¦row44
Vickers | Vickers | 45 | 221-101 | HV 50 | 1 | 490.5 | 20X | 321.971 | 321.549 | 321.76 | 321.8033 | 895.5 | 895.3 | 10.2999 | 1.1504 | 321.004 | 320.734 | 320.869 | 321.2492 | 900.5 | 898.3667 | 11.1392 | 1.2399 |  | 20X | HV 50 | 50 |  | Vickers¦SN=221-101¦scale=HV 50¦F=490.5N¦mag=20X | Vickers¦221-101¦HV 50¦1¦row45
Vickers | Vickers | 46 | 221-101 | HV 50 | 2 | 490.5 | 20X | 321.852 | 322.044 | 321.948 |  | 894.5 |  |  |  | 321.677 | 321.204 | 321.4405 |  | 897.3 |  |  |  |  | 20X | HV 50 | 50 |  | Vickers¦SN=221-101¦scale=HV 50¦F=490.5N¦mag=20X | Vickers¦221-101¦HV 50¦2¦row46
Vickers | Vickers | 47 | 221-101 | HV 50 | 3 | 490.5 | 20X | 321.762 | 321.642 | 321.702 |  | 895.9 |  |  |  | 321.678 | 321.198 | 321.438 |  | 897.3 |  |  |  |  | 20X | HV 50 | 50 |  | Vickers¦SN=221-101¦scale=HV 50¦F=490.5N¦mag=20X | Vickers¦221-101¦HV 50¦3¦row47
Vickers | Vickers | 48 | 293-250 | HV 100 | 1 | 981 | 20X | 1066.623 | 1071.264 | 1068.9435 | 1069.2362 | 162.3 | 162.2 | 7.7835 | 4.7987 | 1066.977 | 1071.568 | 1069.2725 | 1069.524 | 162.2 | 162.1333 | 7.7406 | 4.7742 |  | 20X | HV 100 | 100 |  | Vickers¦SN=293-250¦scale=HV 100¦F=981N¦mag=20X | Vickers¦293-250¦HV 100¦1¦row48
Vickers | Vickers | 49 | 293-250 | HV 100 | 2 | 981 | 20X | 1072.357 | 1072.534 | 1072.4455 |  | 161.2 |  |  |  | 1072.775 | 1072.623 | 1072.699 |  | 161.2 |  |  |  |  | 20X | HV 100 | 100 |  | Vickers¦SN=293-250¦scale=HV 100¦F=981N¦mag=20X | Vickers¦293-250¦HV 100¦2¦row49
Vickers | Vickers | 50 | 293-250 | HV 100 | 3 | 981 | 20X | 1064.814 | 1067.825 | 1066.3195 |  | 163.1 |  |  |  | 1065.147 | 1068.054 | 1066.6005 |  | 163 |  |  |  |  | 20X | HV 100 | 100 |  | Vickers¦SN=293-250¦scale=HV 100¦F=981N¦mag=20X | Vickers¦293-250¦HV 100¦3¦row50
Vickers | Vickers | 51 | 294-367 | HV 100 | 1 | 981 | 20X | 781.723 | 781.85 | 781.7865 | 781.4813 | 303.4 | 303.6667 | 3.4297 | 1.1294 | 781.777 | 782.076 | 781.9265 | 781.7035 | 303.3 | 303.5 | 3.3931 | 1.118 |  | 20X | HV 100 | 100 |  | Vickers¦SN=294-367¦scale=HV 100¦F=981N¦mag=20X | Vickers¦294-367¦HV 100¦1¦row51
Vickers | Vickers | 52 | 294-367 | HV 100 | 2 | 981 | 20X | 780.178 | 782.457 | 781.3175 |  | 303.8 |  |  |  | 781.035 | 782.443 | 781.739 |  | 303.5 |  |  |  |  | 20X | HV 100 | 100 |  | Vickers¦SN=294-367¦scale=HV 100¦F=981N¦mag=20X | Vickers¦294-367¦HV 100¦2¦row52
Vickers | Vickers | 53 | 294-367 | HV 100 | 3 | 981 | 20X | 780.459 | 782.221 | 781.34 |  | 303.8 |  |  |  | 780.912 | 781.978 | 781.445 |  | 303.7 |  |  |  |  | 20X | HV 100 | 100 |  | Vickers¦SN=294-367¦scale=HV 100¦F=981N¦mag=20X | Vickers¦294-367¦HV 100¦3¦row53
Vickers | Vickers | 54 | 294-481 | HV 100 | 1 | 981 | 20X | 605.403 | 606.456 | 605.9295 | 605.7117 | 505.1 | 505.4667 | 7.6895 | 1.5213 | 605.16 | 606.295 | 605.7275 | 605.5448 | 505.4 | 505.7333 | 8.955 | 1.7707 |  | 20X | HV 100 | 100 |  | Vickers¦SN=294-481¦scale=HV 100¦F=981N¦mag=20X | Vickers¦294-481¦HV 100¦1¦row54
Vickers | Vickers | 55 | 294-481 | HV 100 | 2 | 981 | 20X | 605.952 | 606.62 | 606.286 |  | 504.5 |  |  |  | 606.088 | 606.642 | 606.365 |  | 504.4 |  |  |  |  | 20X | HV 100 | 100 |  | Vickers¦SN=294-481¦scale=HV 100¦F=981N¦mag=20X | Vickers¦294-481¦HV 100¦2¦row55
Vickers | Vickers | 56 | 294-481 | HV 100 | 3 | 981 | 20X | 604.529 | 605.31 | 604.9195 |  | 506.8 |  |  |  | 604.142 | 604.942 | 604.542 |  | 507.4 |  |  |  |  | 20X | HV 100 | 100 |  | Vickers¦SN=294-481¦scale=HV 100¦F=981N¦mag=20X | Vickers¦294-481¦HV 100¦3¦row56
Vickers | Vickers | 57 | 216-610 | HV 100 | 1 | 981 | 20X | 511.016 | 511.678 | 511.347 | 511.2327 | 709.2 | 709.5333 | 8.5269 | 1.2018 | 511.311 | 511.273 | 511.292 | 511.089 | 709.4 | 709.9667 | 8.7636 | 1.2344 |  | 20X | HV 100 | 100 |  | Vickers¦SN=216-610¦scale=HV 100¦F=981N¦mag=20X | Vickers¦216-610¦HV 100¦1¦row57
Vickers | Vickers | 58 | 216-610 | HV 100 | 2 | 981 | 20X | 510.923 | 510.777 | 510.85 |  | 710.6 |  |  |  | 510.818 | 510.467 | 510.6425 |  | 711.2 |  |  |  |  | 20X | HV 100 | 100 |  | Vickers¦SN=216-610¦scale=HV 100¦F=981N¦mag=20X | Vickers¦216-610¦HV 100¦2¦row58
Vickers | Vickers | 59 | 216-610 | HV 100 | 3 | 981 | 20X | 511.452 | 511.55 | 511.501 |  | 708.8 |  |  |  | 511.357 | 511.308 | 511.3325 |  | 709.3 |  |  |  |  | 20X | HV 100 | 100 |  | Vickers¦SN=216-610¦scale=HV 100¦F=981N¦mag=20X | Vickers¦216-610¦HV 100¦3¦row59
Vickers | Vickers | 60 | 221-101 | HV 100 | 1 | 981 | 20X | 455.893 | 456.194 | 456.0435 | 456.0675 | 891.7 | 891.6 | 10.0015 | 1.1217 | 455.915 | 456.254 | 456.0845 | 455.9177 | 891.5 | 892.1667 | 10.0253 | 1.1237 |  | 20X | HV 100 | 100 |  | Vickers¦SN=221-101¦scale=HV 100¦F=981N¦mag=20X | Vickers¦221-101¦HV 100¦1¦row60
Vickers | Vickers | 61 | 221-101 | HV 100 | 2 | 981 | 20X | 455.818 | 456.02 | 455.919 |  | 892.2 |  |  |  | 456.009 | 455.49 | 455.7495 |  | 892.8 |  |  |  |  | 20X | HV 100 | 100 |  | Vickers¦SN=221-101¦scale=HV 100¦F=981N¦mag=20X | Vickers¦221-101¦HV 100¦2¦row61
Vickers | Vickers | 62 | 221-101 | HV 100 | 3 | 981 | 20X | 456.126 | 456.354 | 456.24 |  | 890.9 |  |  |  | 455.607 | 456.231 | 455.919 |  | 892.2 |  |  |  |  | 20X | HV 100 | 100 |  | Vickers¦SN=221-101¦scale=HV 100¦F=981N¦mag=20X | Vickers¦221-101¦HV 100¦3¦row62

## Clean_Micro_Vickers
method | original_sheet | excel_row | serial_number | hardness_scale_raw | position | applied_force_n | magnification_raw | manual_d1_um | manual_d2_um | manual_mean_dimension_um | manual_series_mean_dimension_um | manual_hardness | manual_series_mean_hardness | manual_expanded_uncertainty | manual_relative_expanded_uncertainty | automatic_d1_um | automatic_d2_um | automatic_mean_dimension_um | automatic_series_mean_dimension_um | automatic_hardness | automatic_series_mean_hardness | automatic_expanded_uncertainty | automatic_relative_expanded_uncertainty | picture_id | magnification | hardness_scale | hardness_level | ball_diameter_mm | series_id | record_id
Micro-Vickers | Mirco-Vickers | 3 | 710-404 | HV 0,5 | 1 | 0.4905 | 50X | 44.878 | 43.552 | 44.215 | 43.3358 | 47.418 | 49.4187 | 1.0647 | 2.1544 | 44.954 | 43.802 | 44.378 | 43.4723 | 47.07 | 49.1073 | 1.0535 | 2.1453 | 710-404_39.6_HMV_0.05_1_50X | 50X | HV 0.05 | 0.05 |  | Micro-Vickers¦SN=710-404¦scale=HV 0.05¦F=0.4905N¦mag=50X | Micro-Vickers¦710-404¦HV 0.05¦1¦row3
Micro-Vickers | Mirco-Vickers | 4 | 710-404 | HV 0,5 | 2 | 0.4905 | 50X | 43.978 | 43.227 | 43.6025 |  | 48.759 |  |  |  | 43.997 | 43.381 | 43.689 |  | 48.566 |  |  |  |  | 50X | HV 0.05 | 0.05 |  | Micro-Vickers¦SN=710-404¦scale=HV 0.05¦F=0.4905N¦mag=50X | Micro-Vickers¦710-404¦HV 0.05¦2¦row4
Micro-Vickers | Mirco-Vickers | 5 | 710-404 | HV 0,5 | 3 | 0.4905 | 50X | 42.364 | 42.016 | 42.19 |  | 52.079 |  |  |  | 42.393 | 42.307 | 42.35 |  | 51.686 |  |  |  |  | 50X | HV 0.05 | 0.05 |  | Micro-Vickers¦SN=710-404¦scale=HV 0.05¦F=0.4905N¦mag=50X | Micro-Vickers¦710-404¦HV 0.05¦3¦row5
Micro-Vickers | Mirco-Vickers | 6 | 711-081 | HV 0,5 | 1 | 0.4905 | 50X | 28.754 | 28.431 | 28.5925 | 28.7642 | 113.39 | 112.061 | 1.9574 | 1.7467 | 28.748 | 28.501 | 28.6245 | 28.8307 | 113.141 | 111.551 | 1.9569 | 1.7542 | 711-081_102_HMV_0.05_1_50X | 50X | HV 0.05 | 0.05 |  | Micro-Vickers¦SN=711-081¦scale=HV 0.05¦F=0.4905N¦mag=50X | Micro-Vickers¦711-081¦HV 0.05¦1¦row6
Micro-Vickers | Mirco-Vickers | 7 | 711-081 | HV 0,5 | 2 | 0.4905 | 50X | 29.203 | 28.962 | 29.0825 |  | 109.601 |  |  |  | 29.39 | 28.959 | 29.1745 |  | 108.915 |  |  |  |  | 50X | HV 0.05 | 0.05 |  | Micro-Vickers¦SN=711-081¦scale=HV 0.05¦F=0.4905N¦mag=50X | Micro-Vickers¦711-081¦HV 0.05¦2¦row7
Micro-Vickers | Mirco-Vickers | 8 | 711-081 | HV 0,5 | 3 | 0.4905 | 50X | 29.042 | 28.193 | 28.6175 |  | 113.192 |  |  |  | 29.243 | 28.143 | 28.693 |  | 112.597 |  |  |  |  | 50X | HV 0.05 | 0.05 |  | Micro-Vickers¦SN=711-081¦scale=HV 0.05¦F=0.4905N¦mag=50X | Micro-Vickers¦711-081¦HV 0.05¦3¦row8
Micro-Vickers | Mirco-Vickers | 9 | 710-655 | HV 0,5 | 1 | 0.4905 | 50X | 20.888 | 20.796 | 20.842 | 20.8492 | 217.337 | 214.576 | 4.2145 | 1.9641 | 20.921 | 20.887 | 20.904 | 20.8857 | 212.139 | 212.5323 | 4.1871 | 1.9701 | 710-655_197_HMV_0.05_1_50X | 50X | HV 0.05 | 0.05 |  | Micro-Vickers¦SN=710-655¦scale=HV 0.05¦F=0.4905N¦mag=50X | Micro-Vickers¦710-655¦HV 0.05¦1¦row9
Micro-Vickers | Mirco-Vickers | 10 | 710-655 | HV 0,5 | 2 | 0.4905 | 50X | 20.584 | 20.941 | 20.7625 |  | 215.041 |  |  |  | 20.441 | 21.023 | 20.732 |  | 215.674 |  |  |  |  | 50X | HV 0.05 | 0.05 |  | Micro-Vickers¦SN=710-655¦scale=HV 0.05¦F=0.4905N¦mag=50X | Micro-Vickers¦710-655¦HV 0.05¦2¦row10
Micro-Vickers | Mirco-Vickers | 11 | 710-655 | HV 0,5 | 3 | 0.4905 | 50X | 20.701 | 21.185 | 20.943 |  | 211.35 |  |  |  | 20.762 | 21.28 | 21.021 |  | 209.784 |  |  |  |  | 50X | HV 0.05 | 0.05 |  | Micro-Vickers¦SN=710-655¦scale=HV 0.05¦F=0.4905N¦mag=50X | Micro-Vickers¦710-655¦HV 0.05¦3¦row11
Micro-Vickers | Mirco-Vickers | 12 | 711-081 | HV 0,5 | 1 | 4.905 | 50X | 94.577 | 94.173 | 94.375 | 93.8983 | 104.08 | 105.1503 | 1.6828 | 1.6003 | 94.669 | 94.387 | 94.528 | 94.0422 | 103.743 | 104.827 | 1.6047 | 1.5308 | 711-081_102_HMV_0.5_1_50X | 50X | HV 0.5 | 0.5 |  | Micro-Vickers¦SN=711-081¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers¦711-081¦HV 0.5¦1¦row12
Micro-Vickers | Mirco-Vickers | 13 | 711-081 | HV 0,5 | 2 | 4.905 | 50X | 93.058 | 93.169 | 93.1135 |  | 106.919 |  |  |  | 93.436 | 93.271 | 93.3535 |  | 106.371 |  |  |  |  | 50X | HV 0.5 | 0.5 |  | Micro-Vickers¦SN=711-081¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers¦711-081¦HV 0.5¦2¦row13
Micro-Vickers | Mirco-Vickers | 14 | 711-081 | HV 0,5 | 3 | 4.905 | 50X | 94.391 | 94.022 | 94.2065 |  | 104.452 |  |  |  | 94.581 | 93.909 | 94.245 |  | 104.367 |  |  |  |  | 50X | HV 0.5 | 0.5 |  | Micro-Vickers¦SN=711-081¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers¦711-081¦HV 0.5¦3¦row14
Micro-Vickers | Mirco-Vickers | 15 | 706-437 | HV 0,5 | 1 | 4.905 | 50X | 54.053 | 53.972 | 54.0125 | 53.9568 | 318.225 | 318.5757 | 4.4691 | 1.4029 | 53.836 | 53.96 | 53.898 | 54.0013 | 319.106 | 317.8967 | 4.4913 | 1.4128 | 706-437_307_HMV_0.5_1_50X | 50X | HV 0.5 | 0.5 |  | Micro-Vickers¦SN=706-437¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers¦706-437¦HV 0.5¦1¦row15
Micro-Vickers | Mirco-Vickers | 16 | 706-437 | HV 0,5 | 2 | 4.905 | 50X | 53.833 | 53.645 | 53.739 |  | 320.997 |  |  |  | 53.868 | 53.824 | 53.846 |  | 319.722 |  |  |  |  | 50X | HV 0.5 | 0.5 |  | Micro-Vickers¦SN=706-437¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers¦706-437¦HV 0.5¦2¦row16
Micro-Vickers | Mirco-Vickers | 17 | 706-437 | HV 0,5 | 3 | 4.905 | 50X | 54.12 | 54.118 | 54.119 |  | 316.505 |  |  |  | 53.817 | 54.703 | 54.26 |  | 314.862 |  |  |  |  | 50X | HV 0.5 | 0.5 |  | Micro-Vickers¦SN=706-437¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers¦706-437¦HV 0.5¦3¦row17
Micro-Vickers | Mirco-Vickers | 18 | 711-246 | HV 0,5 | 1 | 4.905 | 50X | 42.855 | 42.925 | 42.89 | 42.603 | 503.927 | 510.7747 | 7.7559 | 1.5185 | 42.965 | 43.139 | 43.052 | 42.7943 | 500.142 | 506.258 | 8.0059 | 1.5814 | 711-246_502_HMV_0.5_1_50X | 50X | HV 0.5 | 0.5 |  | Micro-Vickers¦SN=711-246¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers¦711-246¦HV 0.5¦1¦row18
Micro-Vickers | Mirco-Vickers | 19 | 711-246 | HV 0,5 | 2 | 4.905 | 50X | 42.539 | 42.309 | 42.424 |  | 515.058 |  |  |  | 42.483 | 42.216 | 42.3495 |  | 516.86 |  |  |  |  | 50X | HV 0.5 | 0.5 |  | Micro-Vickers¦SN=711-246¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers¦711-246¦HV 0.5¦2¦row19
Micro-Vickers | Mirco-Vickers | 20 | 711-246 | HV 0,5 | 3 | 4.905 | 50X | 42.506 | 42.484 | 42.495 |  | 513.339 |  |  |  | 42.723 | 43.24 | 42.9815 |  | 501.772 |  |  |  |  | 50X | HV 0.5 | 0.5 |  | Micro-Vickers¦SN=711-246¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers¦711-246¦HV 0.5¦3¦row20
Micro-Vickers | Mirco-Vickers | 21 | 710-106 | HV 0,5 | 1 | 4.905 | 50X | 35.089 | 35.34 | 35.2145 | 35.1163 | 747.544 | 751.7683 | 11.9766 | 1.5931 | 35.068 | 34.808 | 34.938 | 35.0033 | 759.423 | 756.691 | 12.3385 | 1.6306 | 710-106_708_HMV_0.5_1_50X | 50X | HV 0.5 | 0.5 |  | Micro-Vickers¦SN=710-106¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers¦710-106¦HV 0.5¦1¦row21
Micro-Vickers | Mirco-Vickers | 22 | 710-106 | HV 0,5 | 2 | 4.905 | 50X | 35.316 | 35.134 | 35.225 |  | 747.098 |  |  |  | 35.396 | 35.236 | 35.316 |  | 743.253 |  |  |  |  | 50X | HV 0.5 | 0.5 |  | Micro-Vickers¦SN=710-106¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers¦710-106¦HV 0.5¦2¦row22
Micro-Vickers | Mirco-Vickers | 23 | 710-106 | HV 0,5 | 3 | 4.905 | 50X | 34.912 | 34.907 | 34.9095 |  | 760.663 |  |  |  | 34.438 | 35.074 | 34.756 |  | 767.397 |  |  |  |  | 50X | HV 0.5 | 0.5 |  | Micro-Vickers¦SN=710-106¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers¦710-106¦HV 0.5¦3¦row23
Micro-Vickers | Mirco-Vickers | 24 | 709-293 | HV 0,5 | 1 | 4.905 | 50X | 30.615 | 30.99 | 30.8025 | 30.7273 | 977.03 | 981.8373 | 16.2799 | 1.6581 | 30.856 | 31.007 | 30.9315 | 30.7503 | 968.929 | 980.475 | 16.5296 | 1.6859 | 709-293_904_HMV_0.5_1_50X | 50X | HV 0.5 | 0.5 |  | Micro-Vickers¦SN=709-293¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers¦709-293¦HV 0.5¦1¦row24
Micro-Vickers | Mirco-Vickers | 25 | 709-293 | HV 0,5 | 2 | 4.905 | 50X | 30.634 | 30.901 | 30.7675 |  | 979.254 |  |  |  | 30.597 | 31.047 | 30.822 |  | 975.794 |  |  |  |  | 50X | HV 0.5 | 0.5 |  | Micro-Vickers¦SN=709-293¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers¦709-293¦HV 0.5¦2¦row25
Micro-Vickers | Mirco-Vickers | 26 | 709-293 | HV 0,5 | 3 | 4.905 | 50X | 30.769 | 30.455 | 30.612 |  | 989.228 |  |  |  | 30.544 | 30.451 | 30.4975 |  | 996.702 |  |  |  |  | 50X | HV 0.5 | 0.5 |  | Micro-Vickers¦SN=709-293¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers¦709-293¦HV 0.5¦3¦row26
Micro-Vickers | Mirco-Vickers | 27 | 711-081 | HV 1 | 1 | 9.81 | 20X | 133.989 | 133.847 | 133.918 | 133.8632 | 103.379 | 103.4753 | 1.3348 | 1.29 | 134.227 | 134.149 | 134.188 | 133.9907 | 102.963 | 103.2693 | 1.3956 | 1.3514 | 711-081_102_HMV_1_1_20X | 20X | HV 1 | 1 |  | Micro-Vickers¦SN=711-081¦scale=HV 1¦F=9.81N¦mag=20X | Micro-Vickers¦711-081¦HV 1¦1¦row27
Micro-Vickers | Mirco-Vickers | 28 | 711-081 | HV 1 | 2 | 9.81 | 20X | 133.772 | 134.662 | 134.217 |  | 102.949 |  |  |  | 133.884 | 134.768 | 134.326 |  | 102.752 |  |  |  |  | 20X | HV 1 | 1 |  | Micro-Vickers¦SN=711-081¦scale=HV 1¦F=9.81N¦mag=20X | Micro-Vickers¦711-081¦HV 1¦2¦row28
Micro-Vickers | Mirco-Vickers | 29 | 711-081 | HV 1 | 3 | 9.81 | 20X | 133.896 | 133.013 | 133.4545 |  | 104.098 |  |  |  | 133.959 | 132.957 | 133.458 |  | 104.093 |  |  |  |  | 20X | HV 1 | 1 |  | Micro-Vickers¦SN=711-081¦scale=HV 1¦F=9.81N¦mag=20X | Micro-Vickers¦711-081¦HV 1¦3¦row29
Micro-Vickers | Mirco-Vickers | 30 | 706-437 | HV 1 | 1 | 9.81 | 50X | 76.986 | 78.473 | 77.7295 | 76.676 | 306.858 | 315.334 | 6.3542 | 2.0151 | 76.23 | 78.803 | 77.5165 | 76.3145 | 308.543 | 318.487 | 6.8887 | 2.1629 | 706-437_307_HMV_1_1_50X | 50X | HV 1 | 1 |  | Micro-Vickers¦SN=706-437¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers¦706-437¦HV 1¦1¦row30
Micro-Vickers | Mirco-Vickers | 31 | 706-437 | HV 1 | 2 | 9.81 | 50X | 76.099 | 75.238 | 75.6685 |  | 323.417 |  |  |  | 75.404 | 75.034 | 75.219 |  | 327.684 |  |  |  | 706-437_307_HMV_1_2_50X | 50X | HV 1 | 1 |  | Micro-Vickers¦SN=706-437¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers¦706-437¦HV 1¦2¦row31
Micro-Vickers | Mirco-Vickers | 32 | 706-437 | HV 1 | 3 | 9.81 | 50X | 76.978 | 76.282 | 76.63 |  | 315.727 |  |  |  | 76.603 | 75.813 | 76.208 |  | 319.234 |  |  |  |  | 50X | HV 1 | 1 |  | Micro-Vickers¦SN=706-437¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers¦706-437¦HV 1¦3¦row32
Micro-Vickers | Mirco-Vickers | 33 | 711-246 | HV 1 | 1 | 9.81 | 50X | 61.233 | 60.571 | 60.902 | 60.518 | 499.858 | 506.2687 | 7.449 | 1.4713 | 61.339 | 60.506 | 60.9225 | 60.277 | 499.513 | 510.365 | 8.0699 | 1.5812 | 711-246_502_HMV_1_1_50X | 50X | HV 1 | 1 |  | Micro-Vickers¦SN=711-246¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers¦711-246¦HV 1¦1¦row33
Micro-Vickers | Mirco-Vickers | 34 | 711-246 | HV 1 | 2 | 9.81 | 50X | 60.38 | 60.759 | 60.5695 |  | 505.361 |  |  |  | 59.748 | 60.131 | 59.9395 |  | 516.049 |  |  |  |  | 50X | HV 1 | 1 |  | Micro-Vickers¦SN=711-246¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers¦711-246¦HV 1¦2¦row34
Micro-Vickers | Mirco-Vickers | 35 | 711-246 | HV 1 | 3 | 9.81 | 50X | 60.709 | 59.456 | 60.0825 |  | 513.587 |  |  |  | 60.284 | 59.654 | 59.969 |  | 515.533 |  |  |  | 711-246_502_HMV_1_3_50X | 50X | HV 1 | 1 |  | Micro-Vickers¦SN=711-246¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers¦711-246¦HV 1¦3¦row35
Micro-Vickers | Mirco-Vickers | 36 | 710-106 | HV 1 | 1 | 9.81 | 50X | 50.597 | 50.739 | 50.668 | 50.3497 | 722.175 | 731.3917 | 10.7947 | 1.4759 | 50.109 | 50.502 | 50.3055 | 50.0255 | 732.635 | 740.9113 | 10.9011 | 1.4713 | 710-106_708_HMV_1_1_50X | 50X | HV 1 | 1 |  | Micro-Vickers¦SN=710-106¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers¦710-106¦HV 1¦1¦row36
Micro-Vickers | Mirco-Vickers | 37 | 710-106 | HV 1 | 2 | 9.81 | 50X | 50.4 | 50.273 | 50.3365 |  | 731.718 |  |  |  | 50.242 | 49.873 | 50.0575 |  | 739.912 |  |  |  |  | 50X | HV 1 | 1 |  | Micro-Vickers¦SN=710-106¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers¦710-106¦HV 1¦2¦row37
Micro-Vickers | Mirco-Vickers | 38 | 710-106 | HV 1 | 3 | 9.81 | 50X | 49.855 | 50.234 | 50.0445 |  | 740.282 |  |  |  | 49.512 | 49.915 | 49.7135 |  | 750.187 |  |  |  |  | 50X | HV 1 | 1 |  | Micro-Vickers¦SN=710-106¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers¦710-106¦HV 1¦3¦row38
Micro-Vickers | Mirco-Vickers | 39 | 709-293 | HV 1 | 1 | 9.81 | 50X | 45.061 | 45.141 | 45.101 | 45.1108 | 911.46 | 911.0723 | 13.23 | 1.4521 | 43.991 | 44.591 | 44.291 | 44.3858 | 945.102 | 941.148 | 14.2215 | 1.5111 | 709-293_904_HMV_1_1_50X | 50X | HV 1 | 1 |  | Micro-Vickers¦SN=709-293¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers¦709-293¦HV 1¦1¦row39
Micro-Vickers | Mirco-Vickers | 40 | 709-293 | HV 1 | 2 | 9.81 | 50X | 45.165 | 44.858 | 45.0115 |  | 915.088 |  |  |  | 44.133 | 44.203 | 44.168 |  | 950.373 |  |  |  |  | 50X | HV 1 | 1 |  | Micro-Vickers¦SN=709-293¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers¦709-293¦HV 1¦2¦row40
Micro-Vickers | Mirco-Vickers | 41 | 709-293 | HV 1 | 3 | 9.81 | 50X | 45.172 | 45.268 | 45.22 |  | 906.669 |  |  |  | 44.82 | 44.577 | 44.6985 |  | 927.969 |  |  |  |  | 50X | HV 1 | 1 |  | Micro-Vickers¦SN=709-293¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers¦709-293¦HV 1¦3¦row41
Micro-Vickers | Mirco-Vickers | 42 | 711-081 | HV 2 | 1 | 19.62 | 20X | 190.789 | 189.096 | 189.9425 | 189.3157 | 102.777 | 103.472 | 2.637 | 2.5485 | 191.047 | 189.242 | 190.1445 | 189.6583 | 102.558 | 103.097 | 2.556 | 2.4793 | 711-081_102_HMV_2_1_20X | 20X | HV 2 | 2 |  | Micro-Vickers¦SN=711-081¦scale=HV 2¦F=19.62N¦mag=20X | Micro-Vickers¦711-081¦HV 2¦1¦row42
Micro-Vickers | Mirco-Vickers | 43 | 711-081 | HV 2 | 2 | 19.62 | 20X | 188.815 | 186.379 | 187.597 |  | 105.363 |  |  |  | 188.96 | 187.083 | 188.0215 |  | 104.887 |  |  |  |  | 20X | HV 2 | 2 |  | Micro-Vickers¦SN=711-081¦scale=HV 2¦F=19.62N¦mag=20X | Micro-Vickers¦711-081¦HV 2¦2¦row43
Micro-Vickers | Mirco-Vickers | 44 | 711-081 | HV 2 | 3 | 19.62 | 20X | 190.903 | 189.912 | 190.4075 |  | 102.276 |  |  |  | 191.346 | 190.272 | 190.809 |  | 101.846 |  |  |  |  | 20X | HV 2 | 2 |  | Micro-Vickers¦SN=711-081¦scale=HV 2¦F=19.62N¦mag=20X | Micro-Vickers¦711-081¦HV 2¦3¦row44
Micro-Vickers | Mirco-Vickers | 45 | 706-437 | HV 2 | 1 | 19.62 | 20X | 108.2 | 108.943 | 108.5715 | 108.0538 | 314.563 | 317.597 | 4.4175 | 1.3909 | 108.206 | 109.032 | 108.619 | 108.0802 | 314.288 | 317.443 | 4.4086 | 1.3888 | 706-437_307_HMV_2_1_50X | 20X | HV 2 | 2 |  | Micro-Vickers¦SN=706-437¦scale=HV 2¦F=19.62N¦mag=20X | Micro-Vickers¦706-437¦HV 2¦1¦row45
Micro-Vickers | Mirco-Vickers | 46 | 706-437 | HV 2 | 2 | 19.62 | 20X | 107.881 | 108.039 | 107.96 |  | 318.137 |  |  |  | 107.664 | 108.136 | 107.9 |  | 318.491 |  |  |  |  | 20X | HV 2 | 2 |  | Micro-Vickers¦SN=706-437¦scale=HV 2¦F=19.62N¦mag=20X | Micro-Vickers¦706-437¦HV 2¦2¦row46
Micro-Vickers | Mirco-Vickers | 47 | 706-437 | HV 2 | 3 | 19.62 | 20X | 107.52 | 107.74 | 107.63 |  | 320.091 |  |  |  | 107.642 | 107.801 | 107.7215 |  | 319.55 |  |  |  |  | 20X | HV 2 | 2 |  | Micro-Vickers¦SN=706-437¦scale=HV 2¦F=19.62N¦mag=20X | Micro-Vickers¦706-437¦HV 2¦3¦row47
Micro-Vickers | Mirco-Vickers | 48 | 711-246 | HV 2 | 1 | 19.62 | 50X | 84.549 | 84.795 | 84.672 | 84.9267 | 517.202 | 514.112 | 6.5828 | 1.2804 | 84.358 | 84.768 | 84.563 | 84.8642 | 518.536 | 514.842 | 6.6849 | 1.2984 | 711-246_502_HMV_2_1_50X | 50X | HV 2 | 2 |  | Micro-Vickers¦SN=711-246¦scale=HV 2¦F=19.62N¦mag=50X | Micro-Vickers¦711-246¦HV 2¦1¦row48
Micro-Vickers | Mirco-Vickers | 49 | 711-246 | HV 2 | 2 | 19.62 | 50X | 85.104 | 85.004 | 85.054 |  | 512.567 |  |  |  | 85.039 | 84.949 | 84.994 |  | 513.194 |  |  |  |  | 50X | HV 2 | 2 |  | Micro-Vickers¦SN=711-246¦scale=HV 2¦F=19.62N¦mag=50X | Micro-Vickers¦711-246¦HV 2¦2¦row49
Micro-Vickers | Mirco-Vickers | 50 | 711-246 | HV 2 | 3 | 19.62 | 50X | 85.007 | 85.101 | 85.054 |  | 512.567 |  |  |  | 84.98 | 85.091 | 85.0355 |  | 512.796 |  |  |  |  | 50X | HV 2 | 2 |  | Micro-Vickers¦SN=711-246¦scale=HV 2¦F=19.62N¦mag=50X | Micro-Vickers¦711-246¦HV 2¦3¦row50
Micro-Vickers | Mirco-Vickers | 51 | 710-106 | HV 2 | 1 | 19.62 | 50X | 70.591 | 71.012 | 70.8015 | 70.8057 | 739.699 | 739.6133 | 9.5237 | 1.2877 | 70.467 | 71.053 | 70.76 | 70.5585 | 740.567 | 744.817 | 9.8039 | 1.3163 | 710-106_708_HMV_2_1_50X | 50X | HV 2 | 2 |  | Micro-Vickers¦SN=710-106¦scale=HV 2¦F=19.62N¦mag=50X | Micro-Vickers¦710-106¦HV 2¦1¦row51
Micro-Vickers | Mirco-Vickers | 52 | 710-106 | HV 2 | 2 | 19.62 | 50X | 70.983 | 70.784 | 70.8835 |  | 737.988 |  |  |  | 71.025 | 70.049 | 70.537 |  | 745.256 |  |  |  |  | 50X | HV 2 | 2 |  | Micro-Vickers¦SN=710-106¦scale=HV 2¦F=19.62N¦mag=50X | Micro-Vickers¦710-106¦HV 2¦2¦row52
Micro-Vickers | Mirco-Vickers | 53 | 710-106 | HV 2 | 3 | 19.62 | 50X | 70.853 | 70.611 | 70.732 |  | 741.153 |  |  |  | 70.55 | 70.207 | 70.3785 |  | 748.628 |  |  |  |  | 50X | HV 2 | 2 |  | Micro-Vickers¦SN=710-106¦scale=HV 2¦F=19.62N¦mag=50X | Micro-Vickers¦710-106¦HV 2¦3¦row53
Micro-Vickers | Mirco-Vickers | 54 | 709-293 | HV 2 | 1 | 19.62 | 50X | 63.552 | 63.767 | 63.6595 | 63.2582 | 914.984 | 926.7043 | 13.4406 | 1.4504 | 63.73 | 63.764 | 63.747 | 63.1747 | 912.473 | 929.244 | 14.9652 | 1.6105 | 709-293_904_HMV_2_1_50X | 50X | HV 2 | 2 |  | Micro-Vickers¦SN=709-293¦scale=HV 2¦F=19.62N¦mag=50X | Micro-Vickers¦709-293¦HV 2¦1¦row54
Micro-Vickers | Mirco-Vickers | 55 | 709-293 | HV 2 | 2 | 19.62 | 50X | 63.336 | 63.174 | 63.255 |  | 926.723 |  |  |  | 63.35 | 63.133 | 63.2415 |  | 927.104 |  |  |  |  | 50X | HV 2 | 2 |  | Micro-Vickers¦SN=709-293¦scale=HV 2¦F=19.62N¦mag=50X | Micro-Vickers¦709-293¦HV 2¦2¦row55
Micro-Vickers | Mirco-Vickers | 56 | 709-293 | HV 2 | 3 | 19.62 | 50X | 62.592 | 63.128 | 62.86 |  | 938.406 |  |  |  | 62.142 | 62.929 | 62.5355 |  | 948.155 |  |  |  |  | 50X | HV 2 | 2 |  | Micro-Vickers¦SN=709-293¦scale=HV 2¦F=19.62N¦mag=50X | Micro-Vickers¦709-293¦HV 2¦3¦row56

## Clean_Brinell
method | original_sheet | excel_row | serial_number | hardness_scale_raw | position | ball_diameter_mm | applied_force_n | magnification_raw | manual_d1_um | manual_d2_um | manual_mean_dimension_um | manual_series_mean_dimension_um | manual_hardness | manual_series_mean_hardness | manual_expanded_uncertainty | manual_relative_expanded_uncertainty | automatic_d1_um | automatic_d2_um | automatic_mean_dimension_um | automatic_series_mean_dimension_um | automatic_hardness | automatic_series_mean_hardness | automatic_expanded_uncertainty | automatic_relative_expanded_uncertainty | picture_id | magnification | hardness_scale | hardness_level | series_id | record_id
Brinell | Brinell | 3 | 293-250 | HBW 1/5 | 1 | 1 | 49.05 | 20x | 209.976 | 211.922 | 210.949 | 210.4625 | 141.497 | 142.1683 | 3.9649 | 2.7889 | 211.318 | 213.35 | 212.334 | 211.8828 | 139.636 | 140.241 | 3.3499 | 2.3886 | 100_HBW_1_5_1_20X | 20X | HBW 1/5 | 1 | Brinell¦SN=293-250¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | Brinell¦293-250¦HBW 1/5¦1¦row3
Brinell | Brinell | 4 | 293-250 | HBW 1/5 | 2 | 1 | 49.05 | 20x | 211.436 | 211.192 | 211.314 |  | 141.003 |  |  |  | 212.215 | 212.275 | 212.245 |  | 139.754 |  |  |  |  | 20X | HBW 1/5 | 1 | Brinell¦SN=293-250¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | Brinell¦293-250¦HBW 1/5¦2¦row4
Brinell | Brinell | 5 | 293-250 | HBW 1/5 | 3 | 1 | 49.05 | 20x | 209.49 | 208.759 | 209.1245 |  | 144.005 |  |  |  | 211.252 | 210.887 | 211.0695 |  | 141.333 |  |  |  |  | 20X | HBW 1/5 | 1 | Brinell¦SN=293-250¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | Brinell¦293-250¦HBW 1/5¦3¦row5
Brinell | Brinell | 6 | 294-367 | HBW 1/5 | 1 | 1 | 49.05 | 20x | 154.258 | 154.988 | 154.623 | 154.5825 | 264.758 | 264.8983 | 6.8031 | 2.5682 | 155.476 | 155.559 | 155.5175 | 155.1857 | 261.701 | 262.832 | 6.8719 | 2.6145 | 300_HBW_1_5_1_20X | 20X | HBW 1/5 | 1 | Brinell¦SN=294-367¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | Brinell¦294-367¦HBW 1/5¦1¦row6
Brinell | Brinell | 7 | 294-367 | HBW 1/5 | 2 | 1 | 49.05 | 20x | 154.501 | 154.988 | 154.7445 |  | 264.34 |  |  |  | 155.086 | 155.477 | 155.2815 |  | 262.502 |  |  |  |  | 20X | HBW 1/5 | 1 | Brinell¦SN=294-367¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | Brinell¦294-367¦HBW 1/5¦2¦row7
Brinell | Brinell | 8 | 294-367 | HBW 1/5 | 3 | 1 | 49.05 | 20x | 154.745 | 154.015 | 154.38 |  | 265.597 |  |  |  | 155.118 | 154.398 | 154.758 |  | 264.293 |  |  |  |  | 20X | HBW 1/5 | 1 | Brinell¦SN=294-367¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | Brinell¦294-367¦HBW 1/5¦3¦row8
Brinell | Brinell | 9 | 290-736 | HBW 1/5 | 1 | 1 | 49.05 | 20x | 115.815 | 116.059 | 115.937 | 116.9922 | 472.178 | 463.921 | 19.6775 | 4.2416 | 117.761 | 117.033 | 117.397 | 117.9267 | 460.467 | 456.5573 | 19.075 | 4.178 | 600_HBW_1_5_1_20X | 20X | HBW 1/5 | 1 | Brinell¦SN=290-736¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | Brinell¦290-736¦HBW 1/5¦1¦row9
Brinell | Brinell | 10 | 290-736 | HBW 1/5 | 2 | 1 | 49.05 | 20x | 119.224 | 119.222 | 119.223 |  | 446.421 |  |  |  | 120.442 | 119.595 | 120.0185 |  | 440.498 |  |  |  |  | 20X | HBW 1/5 | 1 | Brinell¦SN=290-736¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | Brinell¦290-736¦HBW 1/5¦2¦row10
Brinell | Brinell | 11 | 290-736 | HBW 1/5 | 3 | 1 | 49.05 | 20x | 115.086 | 116.547 | 115.8165 |  | 473.164 |  |  |  | 115.698 | 117.031 | 116.3645 |  | 468.707 |  |  |  |  | 20X | HBW 1/5 | 1 | Brinell¦SN=290-736¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | Brinell¦290-736¦HBW 1/5¦3¦row11
Brinell | Brinell | 12 | 293-250 | HBW 1/30 | 1 | 1 | 294.3 | 20x | 466.565 | 463.586 | 465.0755 | 464.74 | 166.519 | 166.7753 | 2.7303 | 1.6371 | 470.313 | 465.067 | 467.69 | 466.62 | 164.542 | 165.3487 | 3.4346 | 2.0772 |  | 20X | HBW 1/30 | 1 | Brinell¦SN=293-250¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | Brinell¦293-250¦HBW 1/30¦1¦row12
Brinell | Brinell | 13 | 293-250 | HBW 1/30 | 2 | 1 | 294.3 | 20x | 464.745 | 464.963 | 464.854 |  | 166.688 |  |  |  | 465.691 | 466.693 | 466.192 |  | 165.671 |  |  |  |  | 20X | HBW 1/30 | 1 | Brinell¦SN=293-250¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | Brinell¦293-250¦HBW 1/30¦2¦row13
Brinell | Brinell | 14 | 293-250 | HBW 1/30 | 3 | 1 | 294.3 | 20x | 463.435 | 465.146 | 464.2905 |  | 167.119 |  |  |  | 465.563 | 466.393 | 465.978 |  | 165.833 |  |  |  |  | 20X | HBW 1/30 | 1 | Brinell¦SN=293-250¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | Brinell¦293-250¦HBW 1/30¦3¦row14
Brinell | Brinell | 15 | 294-367 | HBW 1/30 | 1 | 1 | 294.3 | 20x | 351.556 | 352.297 | 351.9265 | 350.6645 | 298.638 | 300.8717 | 7.3022 | 2.427 | 352.049 | 353.004 | 352.5265 | 351.1102 | 297.587 | 300.084 | 7.5794 | 2.5258 |  | 20X | HBW 1/30 | 1 | Brinell¦SN=294-367¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | Brinell¦294-367¦HBW 1/30¦1¦row15
Brinell | Brinell | 16 | 294-367 | HBW 1/30 | 2 | 1 | 294.3 | 20x | 350.291 | 350.468 | 350.3795 |  | 301.371 |  |  |  | 350.619 | 350.478 | 350.5485 |  | 301.072 |  |  |  |  | 20X | HBW 1/30 | 1 | Brinell¦SN=294-367¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | Brinell¦294-367¦HBW 1/30¦2¦row16
Brinell | Brinell | 17 | 294-367 | HBW 1/30 | 3 | 1 | 294.3 | 20x | 349.477 | 349.898 | 349.6875 |  | 302.606 |  |  |  | 349.884 | 350.627 | 350.2555 |  | 301.593 |  |  |  |  | 20X | HBW 1/30 | 1 | Brinell¦SN=294-367¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | Brinell¦294-367¦HBW 1/30¦3¦row17
Brinell | Brinell | 18 | 290-736 | HBW 1/30 | 1 | 1 | 294.3 | 20x | 254.522 | 254.952 | 254.737 | 254.9415 | 579.109 | 578.167 | 11.6423 | 2.0137 | 257.941 | 255.113 | 256.527 | 255.4053 | 570.918 | 576.049 | 14.0586 | 2.4405 |  | 20X | HBW 1/30 | 1 | Brinell¦SN=290-736¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | Brinell¦290-736¦HBW 1/30¦1¦row18
Brinell | Brinell | 19 | 290-736 | HBW 1/30 | 2 | 1 | 294.3 | 20x | 254.446 | 255.036 | 254.741 |  | 579.091 |  |  |  | 254.616 | 255.125 | 254.8705 |  | 578.494 |  |  |  |  | 20X | HBW 1/30 | 1 | Brinell¦SN=290-736¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | Brinell¦290-736¦HBW 1/30¦2¦row19
Brinell | Brinell | 20 | 290-736 | HBW 1/30 | 3 | 1 | 294.3 | 20x | 255.452 | 255.241 | 255.3465 |  | 576.301 |  |  |  | 254.809 | 254.828 | 254.8185 |  | 578.735 |  |  |  |  | 20X | HBW 1/30 | 1 | Brinell¦SN=290-736¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | Brinell¦290-736¦HBW 1/30¦3¦row20
Brinell | Brinell | 21 | 812-006 | HBW 10/500 | 1 | 10 | 4905 | 20x | 2201.248 | 2201.576 | 2201.412 | 2203.1337 | 129.794 | 129.5897 | 8.9941 | 6.9405 | 2231.69 | 2244.673 | 2238.1815 | 2234.308 | 125.511 | 125.954 | 10.7219 | 8.5126 |  | 20X | HBW 10/500 | 10 | Brinell¦SN=812-006¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=20X | Brinell¦812-006¦HBW 10/500¦1¦row21
Brinell | Brinell | 22 | 812-006 | HBW 10/500 | 2 | 10 | 4905 | 20x | 2208.315 | 2208.155 | 2208.235 |  | 128.983 |  |  |  | 2239.174 | 2216.757 | 2227.9655 |  | 126.68 |  |  |  |  | 20X | HBW 10/500 | 10 | Brinell¦SN=812-006¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=20X | Brinell¦812-006¦HBW 10/500¦2¦row22
Brinell | Brinell | 23 | 812-006 | HBW 10/500 | 3 | 10 | 4905 | 20x | 2199.922 | 2199.586 | 2199.754 |  | 129.992 |  |  |  | 2231.057 | 2242.497 | 2236.777 |  | 125.671 |  |  |  |  | 20X | HBW 10/500 | 10 | Brinell¦SN=812-006¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=20X | Brinell¦812-006¦HBW 10/500¦3¦row23
Brinell | Brinell | 24 | 810-202 | HBW 10/500 | 1 | 10 | 4905 | 20x | 1687.008 | 1687.041 | 1687.0245 | 1686.2058 | 222.152 | 222.3693 | 4.7157 | 2.1207 | 1710.994 | 1714.631 | 1712.8125 | 1710.5187 | 215.464 | 216.0637 | 34.9003 | 16.1528 |  | 20X | HBW 10/500 | 10 | Brinell¦SN=810-202¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=20X | Brinell¦810-202¦HBW 10/500¦1¦row24
Brinell | Brinell | 25 | 810-202 | HBW 10/500 | 2 | 10 | 4905 | 20x | 1684.493 | 1685.233 | 1684.863 |  | 222.726 |  |  |  | 1677.755 | 1720.223 | 1698.989 |  | 219.011 |  |  |  |  | 20X | HBW 10/500 | 10 | Brinell¦SN=810-202¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=20X | Brinell¦810-202¦HBW 10/500¦2¦row25
Brinell | Brinell | 26 | 810-202 | HBW 10/500 | 3 | 10 | 4905 | 20x | 1687.056 | 1686.404 | 1686.73 |  | 222.23 |  |  |  | 1729.945 | 1709.564 | 1719.7545 |  | 213.716 |  |  |  |  | 20X | HBW 10/500 | 10 | Brinell¦SN=810-202¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=20X | Brinell¦810-202¦HBW 10/500¦3¦row26
Brinell | Brinell | 27 | 8301-50DW | HBW 10/500 | 1 | 10 | 4905 | 10x | 1206.303 | 1204.281 | 1205.292 | 1204.3738 | 436.763 | 437.4327 | 10.3037 | 2.3555 | 1265.915 | 1195.979 | 1230.947 | 1239.9103 | 418.681 | 412.87 | 134.7277 | 32.632 |  | 10X | HBW 10/500 | 10 | Brinell¦SN=8301-50DW¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=10X | Brinell¦8301-50DW¦HBW 10/500¦1¦row27
Brinell | Brinell | 28 | 8301-50DW | HBW 10/500 | 2 | 10 | 4905 | 10x | 1203.107 | 1202.59 | 1202.8485 |  | 438.546 |  |  |  | 1251.314 | 1277.344 | 1264.329 |  | 396.78 |  |  |  |  | 10X | HBW 10/500 | 10 | Brinell¦SN=8301-50DW¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=10X | Brinell¦8301-50DW¦HBW 10/500¦2¦row28
Brinell | Brinell | 29 | 8301-50DW | HBW 10/500 | 3 | 10 | 4905 | 10x | 1205.397 | 1204.565 | 1204.981 |  | 436.989 |  |  |  | 1237.417 | 1211.493 | 1224.455 |  | 423.149 |  |  |  |  | 10X | HBW 10/500 | 10 | Brinell¦SN=8301-50DW¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=10X | Brinell¦8301-50DW¦HBW 10/500¦3¦row29
Brinell | Brinell | 30 | 812-006 | HBW 10/3000 | 1 | 10 | 29430 | 20x | 4832.528 | 4831.871 | 4832.1995 | 4840.7873 | 153.449 | 152.8677 | 21.8225 | 14.2754 | 4861.851 | 4842.237 | 4852.044 | 4861.091 | 152.108 | 151.5053 | 38.8338 | 25.632 |  | 20X | HBW 10/3000 | 10 | Brinell¦SN=812-006¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | Brinell¦812-006¦HBW 10/3000¦1¦row30
Brinell | Brinell | 31 | 812-006 | HBW 10/3000 | 2 | 10 | 29430 | 20x | 4847.967 | 4853.497 | 4850.732 |  | 152.196 |  |  |  | 4868.731 | 4892.226 | 4880.4785 |  | 150.213 |  |  |  |  | 20X | HBW 10/3000 | 10 | Brinell¦SN=812-006¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | Brinell¦812-006¦HBW 10/3000¦2¦row31
Brinell | Brinell | 32 | 812-006 | HBW 10/3000 | 3 | 10 | 29430 | 20x | 4834.603 | 4844.258 | 4839.4305 |  | 152.958 |  |  |  | 4845.574 | 4855.927 | 4850.7505 |  | 152.195 |  |  |  |  | 20X | HBW 10/3000 | 10 | Brinell¦SN=812-006¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | Brinell¦812-006¦HBW 10/3000¦3¦row32
Brinell | Brinell | 33 | 810-202 | HBW 10/3000 | 1 | 10 | 29430 | 20x | 3807.15 | 3805.587 | 3806.3685 | 3804.9168 | 253.796 | 253.998 | 11.9262 | 4.6954 | 3809.201 | 3832.075 | 3820.638 | 3823.7332 | 251.827 | 251.411 | 60.0155 | 23.8715 |  | 20X | HBW 10/3000 | 10 | Brinell¦SN=810-202¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | Brinell¦810-202¦HBW 10/3000¦1¦row33
Brinell | Brinell | 34 | 810-202 | HBW 10/3000 | 2 | 10 | 29430 | 20x | 3800.624 | 3802.309 | 3801.4665 |  | 254.478 |  |  |  | 3838.127 | 3843.264 | 3840.6955 |  | 249.095 |  |  |  |  | 20X | HBW 10/3000 | 10 | Brinell¦SN=810-202¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | Brinell¦810-202¦HBW 10/3000¦2¦row34
Brinell | Brinell | 35 | 810-202 | HBW 10/3000 | 3 | 10 | 29430 | 20x | 3807.1 | 3806.731 | 3806.9155 |  | 253.72 |  |  |  | 3809.927 | 3809.805 | 3809.866 |  | 253.311 |  |  |  |  | 20X | HBW 10/3000 | 10 | Brinell¦SN=810-202¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | Brinell¦810-202¦HBW 10/3000¦3¦row35
Brinell | Brinell | 36 | 8301-50DW | HBW 10/3000 | 1 | 10 | 29430 | 20x | 2655.324 | 2657.826 | 2656.575 | 2649.1663 | 531.68 | 534.7197 | 52.634 | 9.8433 | 2657.666 | 2668.334 | 2663 | 2655.9918 | 529.069 | 531.926 | 61.5446 | 11.5701 |  | 20X | HBW 10/3000 | 10 | Brinell¦SN=8301-50DW¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | Brinell¦8301-50DW¦HBW 10/3000¦1¦row36
Brinell | Brinell | 37 | 8301-50DW | HBW 10/3000 | 2 | 10 | 29430 | 20x | 2642.302 | 2648.291 | 2645.2965 |  | 536.308 |  |  |  | 2663.775 | 2650.206 | 2656.9905 |  | 531.51 |  |  |  |  | 20X | HBW 10/3000 | 10 | Brinell¦SN=8301-50DW¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | Brinell¦8301-50DW¦HBW 10/3000¦2¦row37
Brinell | Brinell | 38 | 8301-50DW | HBW 10/3000 | 3 | 10 | 29430 | 20x | 2642.272 | 2648.983 | 2645.6275 |  | 536.171 |  |  |  | 2643.732 | 2652.238 | 2647.985 |  | 535.199 |  |  |  |  | 20X | HBW 10/3000 | 10 | Brinell¦SN=8301-50DW¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | Brinell¦8301-50DW¦HBW 10/3000¦3¦row38

## Clean_Knoop
method | original_sheet | excel_row | serial_number | hardness_scale_raw | position | applied_force_n | magnification_raw | manual_d_um | manual_hardness | automatic_d_um | automatic_hardness | delta_um | delta_percent | magnification | hardness_scale | hardness_level | ball_diameter_mm | series_id | record_id
Knoop | Knoop | 3 | 711-081 | HK 0.05 | 1 | 0.4905 | 50x | 79.238 | [5 empty] | 50X | HK 0.05 | 0.05 |  | Knoop¦SN=711-081¦scale=HK 0.05¦F=0.4905N¦mag=50X | Knoop¦711-081¦HK 0.05¦1¦row3
Knoop | Knoop | 4 | 711-081 | HK 0.05 | 2 | 0.4905 | 50x | 81.54 | [5 empty] | 50X | HK 0.05 | 0.05 |  | Knoop¦SN=711-081¦scale=HK 0.05¦F=0.4905N¦mag=50X | Knoop¦711-081¦HK 0.05¦2¦row4
Knoop | Knoop | 5 | 711-081 | HK 0.05 | 3 | 0.4905 | 50x | 81.6 | [5 empty] | 50X | HK 0.05 | 0.05 |  | Knoop¦SN=711-081¦scale=HK 0.05¦F=0.4905N¦mag=50X | Knoop¦711-081¦HK 0.05¦3¦row5
Knoop | Knoop | 6 | 711-246 | HK 0.05 | 1 | 0.4905 | 50x | 36.686 | [5 empty] | 50X | HK 0.05 | 0.05 |  | Knoop¦SN=711-246¦scale=HK 0.05¦F=0.4905N¦mag=50X | Knoop¦711-246¦HK 0.05¦1¦row6
Knoop | Knoop | 7 | 711-246 | HK 0.05 | 2 | 0.4905 | 50x | 36.54 | [5 empty] | 50X | HK 0.05 | 0.05 |  | Knoop¦SN=711-246¦scale=HK 0.05¦F=0.4905N¦mag=50X | Knoop¦711-246¦HK 0.05¦2¦row7
Knoop | Knoop | 8 | 711-246 | HK 0.05 | 3 | 0.4905 | 50x | 36.103 | [5 empty] | 50X | HK 0.05 | 0.05 |  | Knoop¦SN=711-246¦scale=HK 0.05¦F=0.4905N¦mag=50X | Knoop¦711-246¦HK 0.05¦3¦row8
Knoop | Knoop | 9 | 709-293 | HK 0.05 | 1 | 0.4905 | 50x | 28.275 | [5 empty] | 50X | HK 0.05 | 0.05 |  | Knoop¦SN=709-293¦scale=HK 0.05¦F=0.4905N¦mag=50X | Knoop¦709-293¦HK 0.05¦1¦row9
Knoop | Knoop | 10 | 709-293 | HK 0.05 | 2 | 0.4905 | 50x | 28.515 | [5 empty] | 50X | HK 0.05 | 0.05 |  | Knoop¦SN=709-293¦scale=HK 0.05¦F=0.4905N¦mag=50X | Knoop¦709-293¦HK 0.05¦2¦row10
Knoop | Knoop | 11 | 709-293 | HK 0.05 | 3 | 0.4905 | 50x | 28.493 | [5 empty] | 50X | HK 0.05 | 0.05 |  | Knoop¦SN=709-293¦scale=HK 0.05¦F=0.4905N¦mag=50X | Knoop¦709-293¦HK 0.05¦3¦row11
Knoop | Knoop | 12 | 711-081 | HK 1 | 1 | 9.81 | 10x | 372.77 | [5 empty] | 10X | HK 1 | 1 |  | Knoop¦SN=711-081¦scale=HK 1¦F=9.81N¦mag=10X | Knoop¦711-081¦HK 1¦1¦row12
Knoop | Knoop | 13 | 711-081 | HK 1 | 2 | 9.81 | 10x | 372.321 | [5 empty] | 10X | HK 1 | 1 |  | Knoop¦SN=711-081¦scale=HK 1¦F=9.81N¦mag=10X | Knoop¦711-081¦HK 1¦2¦row13
Knoop | Knoop | 14 | 711-081 | HK 1 | 3 | 9.81 | 10x | 370.76 | [5 empty] | 10X | HK 1 | 1 |  | Knoop¦SN=711-081¦scale=HK 1¦F=9.81N¦mag=10X | Knoop¦711-081¦HK 1¦3¦row14
Knoop | Knoop | 15 | 711-081 | HK 1 | 1 | 9.81 | 20x | 372.629 | [5 empty] | 20X | HK 1 | 1 |  | Knoop¦SN=711-081¦scale=HK 1¦F=9.81N¦mag=20X | Knoop¦711-081¦HK 1¦1¦row15
Knoop | Knoop | 16 | 711-081 | HK 1 | 2 | 9.81 | 20x | 372.141 | [5 empty] | 20X | HK 1 | 1 |  | Knoop¦SN=711-081¦scale=HK 1¦F=9.81N¦mag=20X | Knoop¦711-081¦HK 1¦2¦row16
Knoop | Knoop | 17 | 711-081 | HK 1 | 3 | 9.81 | 20x | 370.888 | [5 empty] | 20X | HK 1 | 1 |  | Knoop¦SN=711-081¦scale=HK 1¦F=9.81N¦mag=20X | Knoop¦711-081¦HK 1¦3¦row17
Knoop | Knoop | 18 | 711-246 | HK 1 | 1 | 9.81 | 20x | 168.648 | [5 empty] | 20X | HK 1 | 1 |  | Knoop¦SN=711-246¦scale=HK 1¦F=9.81N¦mag=20X | Knoop¦711-246¦HK 1¦1¦row18
Knoop | Knoop | 19 | 711-246 | HK 1 | 2 | 9.81 | 20x | 168.246 | [5 empty] | 20X | HK 1 | 1 |  | Knoop¦SN=711-246¦scale=HK 1¦F=9.81N¦mag=20X | Knoop¦711-246¦HK 1¦2¦row19
Knoop | Knoop | 20 | 711-246 | HK 1 | 3 | 9.81 | 20x | 168.975 | [5 empty] | 20X | HK 1 | 1 |  | Knoop¦SN=711-246¦scale=HK 1¦F=9.81N¦mag=20X | Knoop¦711-246¦HK 1¦3¦row20
Knoop | Knoop | 21 | 709-293 | HK 1 | 1 | 9.81 | 20x | 130.027 | [5 empty] | 20X | HK 1 | 1 |  | Knoop¦SN=709-293¦scale=HK 1¦F=9.81N¦mag=20X | Knoop¦709-293¦HK 1¦1¦row21
Knoop | Knoop | 22 | 709-293 | HK 1 | 2 | 9.81 | 20x | 130.271 | [5 empty] | 20X | HK 1 | 1 |  | Knoop¦SN=709-293¦scale=HK 1¦F=9.81N¦mag=20X | Knoop¦709-293¦HK 1¦2¦row22
Knoop | Knoop | 23 | 709-293 | HK 1 | 3 | 9.81 | 20x | 130.674 | [5 empty] | 20X | HK 1 | 1 |  | Knoop¦SN=709-293¦scale=HK 1¦F=9.81N¦mag=20X | Knoop¦709-293¦HK 1¦3¦row23
Knoop | Knoop | 24 | 711-081 | HK 2 | 1 | 19.62 | 10x | 530.591 | [5 empty] | 10X | HK 2 | 2 |  | Knoop¦SN=711-081¦scale=HK 2¦F=19.62N¦mag=10X | Knoop¦711-081¦HK 2¦1¦row24
Knoop | Knoop | 25 | 711-081 | HK 2 | 2 | 19.62 | 10x | 533.922 | [5 empty] | 10X | HK 2 | 2 |  | Knoop¦SN=711-081¦scale=HK 2¦F=19.62N¦mag=10X | Knoop¦711-081¦HK 2¦2¦row25
Knoop | Knoop | 26 | 711-081 | HK 2 | 3 | 19.62 | 10x | 531.51 | [5 empty] | 10X | HK 2 | 2 |  | Knoop¦SN=711-081¦scale=HK 2¦F=19.62N¦mag=10X | Knoop¦711-081¦HK 2¦3¦row26
Knoop | Knoop | 27 | 711-081 | HK 2 | 1 | 19.62 | 20x | 529.435 | [5 empty] | 20X | HK 2 | 2 |  | Knoop¦SN=711-081¦scale=HK 2¦F=19.62N¦mag=20X | Knoop¦711-081¦HK 2¦1¦row27
Knoop | Knoop | 28 | 711-081 | HK 2 | 2 | 19.62 | 20x | 532.945 | [5 empty] | 20X | HK 2 | 2 |  | Knoop¦SN=711-081¦scale=HK 2¦F=19.62N¦mag=20X | Knoop¦711-081¦HK 2¦2¦row28
Knoop | Knoop | 29 | 711-081 | HK 2 | 3 | 19.62 | 20x | 530.498 | [5 empty] | 20X | HK 2 | 2 |  | Knoop¦SN=711-081¦scale=HK 2¦F=19.62N¦mag=20X | Knoop¦711-081¦HK 2¦3¦row29
Knoop | Knoop | 30 | 711-246 | HK 2 | 1 | 19.62 | 20x | 238.966 | [5 empty] | 20X | HK 2 | 2 |  | Knoop¦SN=711-246¦scale=HK 2¦F=19.62N¦mag=20X | Knoop¦711-246¦HK 2¦1¦row30
Knoop | Knoop | 31 | 711-246 | HK 2 | 2 | 19.62 | 20x | 238.601 | [5 empty] | 20X | HK 2 | 2 |  | Knoop¦SN=711-246¦scale=HK 2¦F=19.62N¦mag=20X | Knoop¦711-246¦HK 2¦2¦row31
Knoop | Knoop | 32 | 711-246 | HK 2 | 3 | 19.62 | 20x | 238.283 | [5 empty] | 20X | HK 2 | 2 |  | Knoop¦SN=711-246¦scale=HK 2¦F=19.62N¦mag=20X | Knoop¦711-246¦HK 2¦3¦row32
Knoop | Knoop | 33 | 709-293 | HK 2 | 1 | 19.62 | 20x | 185.244 | [5 empty] | 20X | HK 2 | 2 |  | Knoop¦SN=709-293¦scale=HK 2¦F=19.62N¦mag=20X | Knoop¦709-293¦HK 2¦1¦row33
Knoop | Knoop | 34 | 709-293 | HK 2 | 2 | 19.62 | 20x | 185.156 | [5 empty] | 20X | HK 2 | 2 |  | Knoop¦SN=709-293¦scale=HK 2¦F=19.62N¦mag=20X | Knoop¦709-293¦HK 2¦2¦row34
Knoop | Knoop | 35 | 709-293 | HK 2 | 3 | 19.62 | 20x | 186.143 | [5 empty] | 20X | HK 2 | 2 |  | Knoop¦SN=709-293¦scale=HK 2¦F=19.62N¦mag=20X | Knoop¦709-293¦HK 2¦3¦row35

## Series_Summary
series_id | method | serial_number | hardness_scale_raw | hardness_scale | hardness_level | applied_force_n | ball_diameter_mm | magnification | n_positions | excel_rows | picture_ids | manual_dimension_mean_from_individual_um | manual_dimension_sd_from_individual_um | manual_hardness_mean_from_individual | manual_hardness_sd_from_individual | manual_series_mean_dimension_um | manual_series_mean_hardness | manual_expanded_uncertainty | manual_relative_expanded_uncertainty | automatic_dimension_mean_from_individual_um | automatic_dimension_sd_from_individual_um | automatic_hardness_mean_from_individual | automatic_hardness_sd_from_individual | automatic_series_mean_dimension_um | automatic_series_mean_hardness | automatic_expanded_uncertainty | automatic_relative_expanded_uncertainty
Brinell¦SN=290-736¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | Brinell | 290-736 | HBW 1/30 | HBW 1/30 | 1 | 294.3 | 1 | 20X | 3 | 18,19,20 |  | 254.9415 | 0.35075 | 578.167 | 1.616 | 254.9415 | 578.167 | 11.6423 | 2.0137 | 255.4053 | 0.97174 | 576.049 | 4.4452 | 255.4053 | 576.049 | 14.0586 | 2.4405
Brinell¦SN=290-736¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | Brinell | 290-736 | HBW 1/5 | HBW 1/5 | 1 | 49.05 | 1 | 20X | 3 | 9,10,11 | 600_HBW_1_5_1_20X | 116.9922 | 1.9329 | 463.921 | 15.1635 | 116.9922 | 463.921 | 19.6775 | 4.2416 | 117.9267 | 1.8837 | 456.5573 | 14.5052 | 117.9267 | 456.5573 | 19.075 | 4.178
Brinell¦SN=293-250¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | Brinell | 293-250 | HBW 1/30 | HBW 1/30 | 1 | 294.3 | 1 | 20X | 3 | 12,13,14 |  | 464.74 | 0.40473 | 166.7753 | 0.30939 | 464.74 | 166.7753 | 2.7303 | 1.6371 | 466.62 | 0.9328 | 165.3487 | 0.70327 | 466.62 | 165.3487 | 3.4346 | 2.0772
Brinell¦SN=293-250¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | Brinell | 293-250 | HBW 1/5 | HBW 1/5 | 1 | 49.05 | 1 | 20X | 3 | 3,4,5 | 100_HBW_1_5_1_20X | 210.4625 | 1.173 | 142.1683 | 1.6097 | 210.4625 | 142.1683 | 3.9649 | 2.7889 | 211.8828 | 0.70577 | 140.241 | 0.94754 | 211.8828 | 140.241 | 3.3499 | 2.3886
Brinell¦SN=294-367¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | Brinell | 294-367 | HBW 1/30 | HBW 1/30 | 1 | 294.3 | 1 | 20X | 3 | 15,16,17 |  | 350.6645 | 1.1464 | 300.8717 | 2.0306 | 350.6645 | 300.8717 | 7.3022 | 2.427 | 351.1102 | 1.2353 | 300.084 | 2.1781 | 351.1102 | 300.084 | 7.5794 | 2.5258
Brinell¦SN=294-367¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | Brinell | 294-367 | HBW 1/5 | HBW 1/5 | 1 | 49.05 | 1 | 20X | 3 | 6,7,8 | 300_HBW_1_5_1_20X | 154.5825 | 0.18559 | 264.8983 | 0.64014 | 154.5825 | 264.8983 | 6.8031 | 2.5682 | 155.1857 | 0.38871 | 262.832 | 1.3271 | 155.1857 | 262.832 | 6.8719 | 2.6145
Brinell¦SN=810-202¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | Brinell | 810-202 | HBW 10/3000 | HBW 10/3000 | 10 | 29430 | 10 | 20X | 3 | 33,34,35 |  | 3804.9168 | 3.0006 | 253.998 | 0.41743 | 3804.9168 | 253.998 | 11.9262 | 4.6954 | 3823.7332 | 15.6461 | 251.411 | 2.1386 | 3823.7332 | 251.411 | 60.0155 | 23.8715
Brinell¦SN=810-202¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=20X | Brinell | 810-202 | HBW 10/500 | HBW 10/500 | 10 | 4905 | 10 | 20X | 3 | 24,25,26 |  | 1686.2058 | 1.1722 | 222.3693 | 0.31133 | 1686.2058 | 222.3693 | 4.7157 | 2.1207 | 1710.5187 | 10.5711 | 216.0637 | 2.698 | 1710.5187 | 216.0637 | 34.9003 | 16.1528
Brinell¦SN=812-006¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | Brinell | 812-006 | HBW 10/3000 | HBW 10/3000 | 10 | 29430 | 10 | 20X | 3 | 30,31,32 |  | 4840.7873 | 9.3405 | 152.8677 | 0.63137 | 4840.7873 | 152.8677 | 21.8225 | 14.2754 | 4861.091 | 16.8025 | 151.5053 | 1.12 | 4861.091 | 151.5053 | 38.8338 | 25.632
Brinell¦SN=812-006¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=20X | Brinell | 812-006 | HBW 10/500 | HBW 10/500 | 10 | 4905 | 10 | 20X | 3 | 21,22,23 |  | 2203.1337 | 4.495 | 129.5897 | 0.53463 | 2203.1337 | 129.5897 | 8.9941 | 6.9405 | 2234.308 | 5.5375 | 125.954 | 0.6338 | 2234.308 | 125.954 | 10.7219 | 8.5126
Brinell¦SN=8301-50DW¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | Brinell | 8301-50DW | HBW 10/3000 | HBW 10/3000 | 10 | 29430 | 10 | 20X | 3 | 36,37,38 |  | 2649.1663 | 6.4182 | 534.7197 | 2.6333 | 2649.1663 | 534.7197 | 52.634 | 9.8433 | 2655.9918 | 7.5572 | 531.926 | 3.0861 | 2655.9918 | 531.926 | 61.5446 | 11.5701
Brinell¦SN=8301-50DW¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=10X | Brinell | 8301-50DW | HBW 10/500 | HBW 10/500 | 10 | 4905 | 10 | 10X | 3 | 27,28,29 |  | 1204.3738 | 1.3301 | 437.4327 | 0.97077 | 1204.3738 | 437.4327 | 10.3037 | 2.3555 | 1239.9103 | 21.3949 | 412.87 | 14.1123 | 1239.9103 | 412.87 | 134.7277 | 32.632
Knoop¦SN=709-293¦scale=HK 0.05¦F=0.4905N¦mag=50X | Knoop | 709-293 | HK 0.05 | HK 0.05 | 0.05 | 0.4905 |  | 50X | 3 | 9,10,11
Knoop¦SN=709-293¦scale=HK 1¦F=9.81N¦mag=20X | Knoop | 709-293 | HK 1 | HK 1 | 1 | 9.81 |  | 20X | 3 | 21,22,23
Knoop¦SN=709-293¦scale=HK 2¦F=19.62N¦mag=20X | Knoop | 709-293 | HK 2 | HK 2 | 2 | 19.62 |  | 20X | 3 | 33,34,35
Knoop¦SN=711-081¦scale=HK 0.05¦F=0.4905N¦mag=50X | Knoop | 711-081 | HK 0.05 | HK 0.05 | 0.05 | 0.4905 |  | 50X | 3 | 3,4,5
Knoop¦SN=711-081¦scale=HK 1¦F=9.81N¦mag=10X | Knoop | 711-081 | HK 1 | HK 1 | 1 | 9.81 |  | 10X | 3 | 12,13,14
Knoop¦SN=711-081¦scale=HK 1¦F=9.81N¦mag=20X | Knoop | 711-081 | HK 1 | HK 1 | 1 | 9.81 |  | 20X | 3 | 15,16,17
Knoop¦SN=711-081¦scale=HK 2¦F=19.62N¦mag=10X | Knoop | 711-081 | HK 2 | HK 2 | 2 | 19.62 |  | 10X | 3 | 24,25,26
Knoop¦SN=711-081¦scale=HK 2¦F=19.62N¦mag=20X | Knoop | 711-081 | HK 2 | HK 2 | 2 | 19.62 |  | 20X | 3 | 27,28,29
Knoop¦SN=711-246¦scale=HK 0.05¦F=0.4905N¦mag=50X | Knoop | 711-246 | HK 0.05 | HK 0.05 | 0.05 | 0.4905 |  | 50X | 3 | 6,7,8
Knoop¦SN=711-246¦scale=HK 1¦F=9.81N¦mag=20X | Knoop | 711-246 | HK 1 | HK 1 | 1 | 9.81 |  | 20X | 3 | 18,19,20
Knoop¦SN=711-246¦scale=HK 2¦F=19.62N¦mag=20X | Knoop | 711-246 | HK 2 | HK 2 | 2 | 19.62 |  | 20X | 3 | 30,31,32
Micro-Vickers¦SN=706-437¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers | 706-437 | HV 0,5 | HV 0.5 | 0.5 | 4.905 |  | 50X | 3 | 15,16,17 | 706-437_307_HMV_0.5_1_50X | 53.9568 | 0.19602 | 318.5757 | 2.2664 | 53.9568 | 318.5757 | 4.4691 | 1.4029 | 54.0013 | 0.22552 | 317.8967 | 2.6461 | 54.0013 | 317.8967 | 4.4913 | 1.4128
Micro-Vickers¦SN=706-437¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers | 706-437 | HV 1 | HV 1 | 1 | 9.81 |  | 50X | 3 | 30,31,32 | 706-437_307_HMV_1_1_50X; 706-437_307_HMV_1_2_50X | 76.676 | 1.0313 | 315.334 | 8.2865 | 76.676 | 315.334 | 6.3542 | 2.0151 | 76.3145 | 1.1524 | 318.487 | 9.5923 | 76.3145 | 318.487 | 6.8887 | 2.1629
Micro-Vickers¦SN=706-437¦scale=HV 2¦F=19.62N¦mag=20X | Micro-Vickers | 706-437 | HV 2 | HV 2 | 2 | 19.62 |  | 20X | 3 | 45,46,47 | 706-437_307_HMV_2_1_50X | 108.0538 | 0.47771 | 317.597 | 2.8033 | 108.0538 | 317.597 | 4.4175 | 1.3909 | 108.0802 | 0.4751 | 317.443 | 2.7831 | 108.0802 | 317.443 | 4.4086 | 1.3888
Micro-Vickers¦SN=709-293¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers | 709-293 | HV 0,5 | HV 0.5 | 0.5 | 4.905 |  | 50X | 3 | 24,25,26 | 709-293_904_HMV_0.5_1_50X | 30.7273 | 0.1014 | 981.8373 | 6.4964 | 30.7273 | 981.8373 | 16.2799 | 1.6581 | 30.7503 | 0.2257 | 980.475 | 14.4661 | 30.7503 | 980.475 | 16.5296 | 1.6859
Micro-Vickers¦SN=709-293¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers | 709-293 | HV 1 | HV 1 | 1 | 9.81 |  | 50X | 3 | 39,40,41 | 709-293_904_HMV_1_1_50X | 45.1108 | 0.1046 | 911.0723 | 4.2229 | 45.1108 | 911.0723 | 13.23 | 1.4521 | 44.3858 | 0.27767 | 941.148 | 11.7137 | 44.3858 | 941.148 | 14.2215 | 1.5111
Micro-Vickers¦SN=709-293¦scale=HV 2¦F=19.62N¦mag=50X | Micro-Vickers | 709-293 | HV 2 | HV 2 | 2 | 19.62 |  | 50X | 3 | 54,55,56 | 709-293_904_HMV_2_1_50X | 63.2582 | 0.39976 | 926.7043 | 11.711 | 63.2582 | 926.7043 | 13.4406 | 1.4504 | 63.1747 | 0.60851 | 929.244 | 17.937 | 63.1747 | 929.244 | 14.9652 | 1.6105
Micro-Vickers¦SN=710-106¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers | 710-106 | HV 0,5 | HV 0.5 | 0.5 | 4.905 |  | 50X | 3 | 21,22,23 | 710-106_708_HMV_0.5_1_50X | 35.1163 | 0.1792 | 751.7683 | 7.7062 | 35.1163 | 751.7683 | 11.9766 | 1.5931 | 35.0033 | 0.28566 | 756.691 | 12.3017 | 35.0033 | 756.691 | 12.3385 | 1.6306
Micro-Vickers¦SN=710-106¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers | 710-106 | HV 1 | HV 1 | 1 | 9.81 |  | 50X | 3 | 36,37,38 | 710-106_708_HMV_1_1_50X | 50.3497 | 0.31196 | 731.3917 | 9.0579 | 50.3497 | 731.3917 | 10.7947 | 1.4759 | 50.0255 | 0.29729 | 740.9113 | 8.8186 | 50.0255 | 740.9113 | 10.9011 | 1.4713
Micro-Vickers¦SN=710-106¦scale=HV 2¦F=19.62N¦mag=50X | Micro-Vickers | 710-106 | HV 2 | HV 2 | 2 | 19.62 |  | 50X | 3 | 51,52,53 | 710-106_708_HMV_2_1_50X | 70.8057 | 0.075836 | 739.6133 | 1.5842 | 70.8057 | 739.6133 | 9.5237 | 1.2877 | 70.5585 | 0.19166 | 744.817 | 4.0484 | 70.5585 | 744.817 | 9.8039 | 1.3163
Micro-Vickers¦SN=710-404¦scale=HV 0.05¦F=0.4905N¦mag=50X | Micro-Vickers | 710-404 | HV 0,5 | HV 0.05 | 0.05 | 0.4905 |  | 50X | 3 | 3,4,5 | 710-404_39.6_HMV_0.05_1_50X | 43.3358 | 1.0385 | 49.4187 | 2.3995 | 43.3358 | 49.4187 | 1.0647 | 2.1544 | 43.4723 | 1.0312 | 49.1073 | 2.3551 | 43.4723 | 49.1073 | 1.0535 | 2.1453
Micro-Vickers¦SN=710-655¦scale=HV 0.05¦F=0.4905N¦mag=50X | Micro-Vickers | 710-655 | HV 0,5 | HV 0.05 | 0.05 | 0.4905 |  | 50X | 3 | 9,10,11 | 710-655_197_HMV_0.05_1_50X | 20.8492 | 0.090463 | 214.576 | 3.0205 | 20.8492 | 214.576 | 4.2145 | 1.9641 | 20.8857 | 0.14537 | 212.5323 | 2.9646 | 20.8857 | 212.5323 | 4.1871 | 1.9701
Micro-Vickers¦SN=711-081¦scale=HV 0.05¦F=0.4905N¦mag=50X | Micro-Vickers | 711-081 | HV 0,5 | HV 0.05 | 0.05 | 0.4905 |  | 50X | 3 | 6,7,8 | 711-081_102_HMV_0.05_1_50X | 28.7642 | 0.27597 | 112.061 | 2.1327 | 28.7642 | 112.061 | 1.9574 | 1.7467 | 28.8307 | 0.29973 | 111.551 | 2.299 | 28.8307 | 111.551 | 1.9569 | 1.7542
Micro-Vickers¦SN=711-081¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers | 711-081 | HV 0,5 | HV 0.5 | 0.5 | 4.905 |  | 50X | 3 | 12,13,14 | 711-081_102_HMV_0.5_1_50X | 93.8983 | 0.68489 | 105.1503 | 1.543 | 93.8983 | 105.1503 | 1.6828 | 1.6003 | 94.0422 | 0.61296 | 104.827 | 1.3731 | 94.0422 | 104.827 | 1.6047 | 1.5308
Micro-Vickers¦SN=711-081¦scale=HV 1¦F=9.81N¦mag=20X | Micro-Vickers | 711-081 | HV 1 | HV 1 | 1 | 9.81 |  | 20X | 3 | 27,28,29 | 711-081_102_HMV_1_1_20X | 133.8632 | 0.3842 | 103.4753 | 0.58053 | 133.8632 | 103.4753 | 1.3348 | 1.29 | 133.9907 | 0.46643 | 103.2693 | 0.72108 | 133.9907 | 103.2693 | 1.3956 | 1.3514
Micro-Vickers¦SN=711-081¦scale=HV 2¦F=19.62N¦mag=20X | Micro-Vickers | 711-081 | HV 2 | HV 2 | 2 | 19.62 |  | 20X | 3 | 42,43,44 | 711-081_102_HMV_2_1_20X | 189.3157 | 1.5065 | 103.472 | 1.6567 | 189.3157 | 103.472 | 2.637 | 2.5485 | 189.6583 | 1.456 | 103.097 | 1.5905 | 189.6583 | 103.097 | 2.556 | 2.4793
Micro-Vickers¦SN=711-246¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers | 711-246 | HV 0,5 | HV 0.5 | 0.5 | 4.905 |  | 50X | 3 | 18,19,20 | 711-246_502_HMV_0.5_1_50X | 42.603 | 0.25107 | 510.7747 | 5.9922 | 42.603 | 510.7747 | 7.7559 | 1.5185 | 42.7943 | 0.38685 | 506.258 | 9.2177 | 42.7943 | 506.258 | 8.0059 | 1.5814
Micro-Vickers¦SN=711-246¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers | 711-246 | HV 1 | HV 1 | 1 | 9.81 |  | 50X | 3 | 33,34,35 | 711-246_502_HMV_1_1_50X; 711-246_502_HMV_1_3_50X | 60.518 | 0.41217 | 506.2687 | 6.9094 | 60.518 | 506.2687 | 7.449 | 1.4713 | 60.277 | 0.55921 | 510.365 | 9.4016 | 60.277 | 510.365 | 8.0699 | 1.5812
Micro-Vickers¦SN=711-246¦scale=HV 2¦F=19.62N¦mag=50X | Micro-Vickers | 711-246 | HV 2 | HV 2 | 2 | 19.62 |  | 50X | 3 | 48,49,50 | 711-246_502_HMV_2_1_50X | 84.9267 | 0.22055 | 514.112 | 2.676 | 84.9267 | 514.112 | 6.5828 | 1.2804 | 84.8642 | 0.26164 | 514.842 | 3.2053 | 84.8642 | 514.842 | 6.6849 | 1.2984
Vickers¦SN=216-610¦scale=HV 100¦F=981N¦mag=20X | Vickers | 216-610 | HV 100 | HV 100 | 100 | 981 |  | 20X | 3 | 57,58,59 |  | 511.2327 | 0.34023 | 709.5333 | 0.94516 | 511.2327 | 709.5333 | 8.5269 | 1.2018 | 511.089 | 0.38721 | 709.9667 | 1.0693 | 511.089 | 709.9667 | 8.7636 | 1.2344
Vickers¦SN=216-610¦scale=HV 20¦F=196.2N¦mag=20X | Vickers | 216-610 | HV 20 | HV 20 | 20 | 196.2 |  | 20X | 3 | 27,28,29 | 216-610_701_HV_20_1_20X | 227.4537 | 0.39908 | 716.7313 | 2.5185 | 227.4537 | 716.7313 | 9.5727 | 1.3356 | 227.1698 | 0.48427 | 717.44 | 3.0217 | 227.1698 | 717.44 | 10.0422 | 1.3997
Vickers¦SN=216-610¦scale=HV 3¦F=29.43N¦mag=50X | Vickers | 216-610 | HV 3 | HV 3 | 3 | 29.43 |  | 50X | 3 | 12,13,14 | 216-610_701_HV_3_1_50X | 87.4893 | 0.32459 | 726.6623 | 5.3945 | 87.4893 | 726.6623 | 11.4188 | 1.5714 | 87.2083 | 0.83125 | 731.3793 | 13.7487 | 87.2083 | 731.3793 | 14.3219 | 1.9582
Vickers¦SN=216-610¦scale=HV 50¦F=490.5N¦mag=20X | Vickers | 216-610 | HV 50 | HV 50 | 50 | 490.5 |  | 20X | 3 | 42,43,44 |  | 360.7707 | 0.39373 | 712.3667 | 1.5503 | 360.7707 | 712.3667 | 9.0466 | 1.2699 | 360.494 | 0.26377 | 713.4667 | 1.0599 | 360.494 | 713.4667 | 8.4852 | 1.1893
Vickers¦SN=221-101¦scale=HV 100¦F=981N¦mag=20X | Vickers | 221-101 | HV 100 | HV 100 | 100 | 981 |  | 20X | 3 | 60,61,62 |  | 456.0675 | 0.16184 | 891.6 | 0.65574 | 456.0675 | 891.6 | 10.0015 | 1.1217 | 455.9177 | 0.1675 | 892.1667 | 0.65064 | 455.9177 | 892.1667 | 10.0253 | 1.1237
Vickers¦SN=221-101¦scale=HV 20¦F=196.2N¦mag=20X | Vickers | 221-101 | HV 20 | HV 20 | 20 | 196.2 |  | 20X | 3 | 30,31,32 | 221-101_898_HV_20_1_20X | 202.8217 | 0.20328 | 901.3887 | 1.8054 | 202.8217 | 901.3887 | 11.2763 | 1.251 | 202.566 | 0.37204 | 903.6693 | 3.3246 | 202.566 | 903.6693 | 12.0942 | 1.3383
Vickers¦SN=221-101¦scale=HV 3¦F=29.43N¦mag=50X | Vickers | 221-101 | HV 3 | HV 3 | 3 | 29.43 |  | 50X | 3 | 15,16,17 | 221-101_898_HV_3_1_50X | 78.1743 | 0.070671 | 910.1297 | 1.6456 | 78.1743 | 910.1297 | 14.1415 | 1.5538 | 77.9745 | 0.42186 | 914.8567 | 9.9076 | 77.9745 | 914.8567 | 15.3644 | 1.6794
Vickers¦SN=221-101¦scale=HV 50¦F=490.5N¦mag=20X | Vickers | 221-101 | HV 50 | HV 50 | 50 | 490.5 |  | 20X | 3 | 45,46,47 |  | 321.8033 | 0.1286 | 895.3 | 0.72111 | 321.8033 | 895.3 | 10.2999 | 1.1504 | 321.2492 | 0.32924 | 898.3667 | 1.8475 | 321.2492 | 898.3667 | 11.1392 | 1.2399
Vickers¦SN=293-250¦scale=HV 100¦F=981N¦mag=20X | Vickers | 293-250 | HV 100 | HV 100 | 100 | 981 |  | 20X | 3 | 48,49,50 |  | 1069.2362 | 3.0735 | 162.2 | 0.95394 | 1069.2362 | 162.2 | 7.7835 | 4.7987 | 1069.524 | 3.057 | 162.1333 | 0.90185 | 1069.524 | 162.1333 | 7.7406 | 4.7742
Vickers¦SN=293-250¦scale=HV 20¦F=196.2N¦mag=20X | Vickers | 293-250 | HV 20 | HV 20 | 20 | 196.2 |  | 20X | 3 | 18,19,20 |  | 479.2792 | 1.8649 | 161.4267 | 1.2526 | 479.2792 | 161.4267 | 4.9143 | 3.0443 | 479.3273 | 2.4103 | 161.3973 | 1.6166 | 479.3273 | 161.3973 | 6.1848 | 3.832
Vickers¦SN=293-250¦scale=HV 3¦F=29.43N¦mag=20X | Vickers | 293-250 | HV 3 | HV 3 | 3 | 29.43 |  | 20X | 3 | 3,4,5 | 293-250_160_HV_3_1_20X | 184.5562 | 0.61668 | 163.5923 | 0.71999 | 184.5562 | 163.5923 | 2.5363 | 1.5504 | 184.5995 | 0.20798 | 163.2407 | 0.40581 | 184.5995 | 163.2407 | 2.0781 | 1.273
Vickers¦SN=293-250¦scale=HV 50¦F=490.5N¦mag=20X | Vickers | 293-250 | HV 50 | HV 50 | 50 | 490.5 |  | 20X | 3 | 33,34,35 |  | 756.2575 | 2.734 | 162.1 | 1.1533 | 756.2575 | 162.1 | 6.9689 | 4.2991 | 756.645 | 2.8734 | 162.0333 | 1.2503 | 756.645 | 162.0333 | 7.3001 | 4.5053
Vickers¦SN=294-367¦scale=HV 100¦F=981N¦mag=20X | Vickers | 294-367 | HV 100 | HV 100 | 100 | 981 |  | 20X | 3 | 51,52,53 |  | 781.4813 | 0.26452 | 303.6667 | 0.23094 | 781.4813 | 303.6667 | 3.4297 | 1.1294 | 781.7035 | 0.24271 | 303.5 | 0.2 | 781.7035 | 303.5 | 3.3931 | 1.118
Vickers¦SN=294-367¦scale=HV 20¦F=196.2N¦mag=20X | Vickers | 294-367 | HV 20 | HV 20 | 20 | 196.2 |  | 20X | 3 | 21,22,23 |  | 349.196 | 0.35864 | 304.09 | 0.62488 | 349.196 | 304.09 | 3.7984 | 1.2491 | 349.4752 | 0.36651 | 303.6047 | 0.6368 | 349.4752 | 303.6047 | 3.8082 | 1.2543
Vickers¦SN=294-367¦scale=HV 3¦F=29.43N¦mag=20X | Vickers | 294-367 | HV 3 | HV 3 | 3 | 29.43 |  | 20X | 3 | 6,7,8 | 294-367_300_HV_3_1_20X | 135.062 | 0.2847 | 304.9003 | 1.2964 | 135.062 | 304.9003 | 4.2319 | 1.388 | 135.0908 | 0.33266 | 304.4983 | 1.5102 | 135.0908 | 304.4983 | 4.3009 | 1.4124
Vickers¦SN=294-367¦scale=HV 50¦F=490.5N¦mag=20X | Vickers | 294-367 | HV 50 | HV 50 | 50 | 490.5 |  | 20X | 3 | 36,37,38 |  | 551.6933 | 0.43959 | 304.6333 | 0.47258 | 551.6933 | 304.6333 | 3.8664 | 1.2692 | 551.7327 | 0.4336 | 304.5667 | 0.46188 | 551.7327 | 304.5667 | 3.851 | 1.2644
Vickers¦SN=294-481¦scale=HV 100¦F=981N¦mag=20X | Vickers | 294-481 | HV 100 | HV 100 | 100 | 981 |  | 20X | 3 | 54,55,56 |  | 605.7117 | 0.70882 | 505.4667 | 1.193 | 605.7117 | 505.4667 | 7.6895 | 1.5213 | 605.5448 | 0.92513 | 505.7333 | 1.5275 | 605.5448 | 505.7333 | 8.955 | 1.7707
Vickers¦SN=294-481¦scale=HV 20¦F=196.2N¦mag=20X | Vickers | 294-481 | HV 20 | HV 20 | 20 | 196.2 |  | 20X | 3 | 24,25,26 | 294_481_503_HV_20_1_20X | 270.0955 | 0.35338 | 508.284 | 1.3284 | 270.0955 | 508.284 | 6.497 | 1.2782 | 269.9183 | 0.2894 | 508.951 | 1.0906 | 269.9183 | 508.951 | 6.313 | 1.2404
Vickers¦SN=294-481¦scale=HV 3¦F=29.43N¦mag=20X | Vickers | 294-481 | HV 3 | HV 3 | 3 | 29.43 |  | 20X | 3 | 9,10,11 | 294_481_503_HV_3_1_20X | 103.873 | 0.17282 | 515.499 | 1.7126 | 103.873 | 515.499 | 7.4144 | 1.4383 | 104.2282 | 0.22259 | 511.9953 | 2.1799 | 104.2282 | 511.9953 | 7.4379 | 1.4527
Vickers¦SN=294-481¦scale=HV 50¦F=490.5N¦mag=20X | Vickers | 294-481 | HV 50 | HV 50 | 50 | 490.5 |  | 20X | 3 | 39,40,41 |  | 426.5372 | 0.19549 | 509.6 | 0.45826 | 426.5372 | 509.6 | 5.8116 | 1.1404 | 426.8152 | 0.36082 | 508.9333 | 0.86217 | 426.8152 | 508.9333 | 6.2624 | 1.2305

## Paired_Summary
method | n | mean_difference | median_difference | sd_difference | mad_difference | iqr_difference | rmse | mae | symmetric_mean_relative_difference_percent | bias_ci_low | bias_ci_high | cohens_dz | paired_t_p | wilcoxon_p
Brinell | 36 | 11.8931 | 2.486 | 15.2696 | 3.4645 | 18.8892 | 19.1867 | 11.9224 | 0.78932 | 6.7266 | 17.0595 | 0.77887 | 4.2873e-05 | 2.0373e-10
Micro-Vickers | 54 | -0.05662 | 0.012 | 0.27813 | 0.18755 | 0.30137 | 0.2813 | 0.2051 | -0.12486 | -0.13254 | 0.019296 | -0.20357 | 0.1406 | 0.48823
Vickers | 60 | -0.025958 | -0.0015 | 0.36702 | 0.37992 | 0.503 | 0.36488 | 0.30097 | -0.029036 | -0.12077 | 0.068853 | -0.070727 | 0.58586 | 0.68556

## Paired_By_Force
method | applied_force_n | magnification | n | mean_difference | median_difference | sd_difference | mad_difference | iqr_difference | rmse | mae | symmetric_mean_relative_difference_percent | bias_ci_low | bias_ci_high | cohens_dz | paired_t_p | wilcoxon_p
Brinell | 49.05 | 20X | 9 | 0.986 | 0.8945 | 0.5141 | 0.53003 | 0.837 | 1.0987 | 0.986 | 0.61956 | 0.59083 | 1.3812 | 1.9179 | 0.00042714 | 0.0039062
Brinell | 294.3 | 20X | 9 | 0.92983 | 0.6 | 0.99392 | 1.0942 | 1.5185 | 1.3201 | 1.0472 | 0.2373 | 0.16584 | 1.6938 | 0.93552 | 0.022964 | 0.019531
Brinell | 4905 | 10X | 3 | 35.5365 | 25.655 | 22.6797 | 9.164 | 21.0033 | 40.0719 | 35.5365 | 2.8977 | -20.803 | 91.876 | 1.5669 | 0.11318 | 0.25
Brinell | 4905 | 20X | 6 | 27.7436 | 29.4062 | 9.4756 | 11.1047 | 14.5884 | 29.0608 | 27.7436 | 1.4176 | 17.7995 | 37.6876 | 2.9279 | 0.00081976 | 0.03125
Brinell | 29430 | 20X | 9 | 15.3152 | 11.694 | 12.4028 | 12.0839 | 13.4195 | 19.2689 | 15.3152 | 0.38943 | 5.7815 | 24.8488 | 1.2348 | 0.0060041 | 0.0039062
Micro-Vickers | 0.4905 | 50X | 9 | 0.079833 | 0.078 | 0.059516 | 0.023722 | 0.03 | 0.097581 | 0.086611 | 0.23974 | 0.034085 | 0.12558 | 1.3414 | 0.0038192 | 0.0078125
Micro-Vickers | 4.905 | 50X | 15 | 0.057933 | 0.091 | 0.18633 | 0.10526 | 0.2415 | 0.18911 | 0.15573 | 0.086396 | -0.045255 | 0.16112 | 0.31091 | 0.24849 | 0.3028
Micro-Vickers | 9.81 | 20X | 3 | 0.1275 | 0.109 | 0.13421 | 0.15641 | 0.13325 | 0.16812 | 0.1275 | 0.095071 | -0.2059 | 0.4609 | 0.95001 | 0.24161 | 0.25
Micro-Vickers | 9.81 | 50X | 12 | -0.41292 | -0.39225 | 0.26078 | 0.22869 | 0.28613 | 0.48253 | 0.41633 | -0.78539 | -0.57861 | -0.24723 | -1.5834 | 0.00019049 | 0.00097656
Micro-Vickers | 19.62 | 20X | 6 | 0.1845 | 0.14675 | 0.19606 | 0.22684 | 0.29312 | 0.25705 | 0.2045 | 0.10268 | -0.02125 | 0.39025 | 0.94105 | 0.069331 | 0.09375
Micro-Vickers | 19.62 | 50X | 9 | -0.13106 | -0.06 | 0.16626 | 0.072647 | 0.306 | 0.20432 | 0.1505 | -0.18582 | -0.25885 | -0.0032598 | -0.78828 | 0.045613 | 0.039062
Vickers | 29.43 | 20X | 9 | 0.14244 | 0.1375 | 0.27816 | 0.21053 | 0.374 | 0.29844 | 0.24178 | 0.12879 | -0.071366 | 0.35625 | 0.5121 | 0.16302 | 0.25
Vickers | 29.43 | 50X | 6 | -0.24042 | -0.38125 | 0.45945 | 0.40957 | 0.76313 | 0.48344 | 0.45208 | -0.29058 | -0.72258 | 0.24175 | -0.52327 | 0.25615 | 0.3125
Vickers | 196.2 | 20X | 15 | -0.077867 | -0.0155 | 0.40625 | 0.35582 | 0.399 | 0.40013 | 0.3142 | -0.04542 | -0.30284 | 0.14711 | -0.19167 | 0.47016 | 0.45428
Vickers | 490.5 | 20X | 15 | -0.0252 | -0.141 | 0.42392 | 0.44626 | 0.53025 | 0.41032 | 0.35373 | -0.025126 | -0.25996 | 0.20956 | -0.059446 | 0.82124 | 0.84692
Vickers | 981 | 20X | 15 | 0.0099333 | 0.041 | 0.24723 | 0.31505 | 0.3825 | 0.23906 | 0.21007 | -0.0066408 | -0.12698 | 0.14685 | 0.040178 | 0.87856 | 0.93408

## Bland_Altman
method | representation | n | bias | sd_diff | lower_loa | upper_loa | n_ci_low | n_ci_high | bias_ci_low | bias_ci_high | sd_diff_ci_low | sd_diff_ci_high | lower_loa_ci_low | lower_loa_ci_high | upper_loa_ci_low | upper_loa_ci_high | proportional_bias_slope | proportional_bias_p | absdiff_pairmean_spearman | measurand
Brinell | absolute | 36 | 11.8931 | 15.2696 | -18.0353 | 41.8214 | 36 | 36 | 5.1925 | 18.9633 | 9.3859 | 18.5686 | -22.3231 | -10.561 | 24.6396 | 53.9446 | 0.0046659 | 0.0036346 | 0.74903 | dimension
Brinell | relative_percent | 36 | 0.78932 | 0.92421 | -1.0221 | 2.6008 | 36 | 36 | 0.42593 | 1.25 | 0.34599 | 1.3639 | -1.5796 | -0.2183 | 1.1228 | 3.8244 | 1.9421e-06 | 0.985 | 0.052252 | dimension
Brinell | log_ratio | 36 | 0.0078936 | 0.0092435 | -0.010224 | 0.026011 | 36 | 36 | 0.0042594 | 0.012501 | 0.0034599 | 0.013642 | -0.0158 | -0.0021832 | 0.011228 | 0.03825 | 0.0012914 | 0.29423 | 0.052252 | dimension
Micro-Vickers | absolute | 54 | -0.05662 | 0.27813 | -0.60176 | 0.48852 | 54 | 54 | -0.18134 | 0.049445 | 0.18514 | 0.35582 | -0.8487 | -0.34166 | 0.33752 | 0.6194 | 0.0021505 | 0.018123 | 0.08161 | dimension
Micro-Vickers | relative_percent | 54 | -0.12486 | 0.53479 | -1.173 | 0.92333 | 54 | 54 | -0.37886 | 0.070702 | 0.31807 | 0.73838 | -1.7626 | -0.58735 | 0.59278 | 1.1962 | 0.0017079 | 0.33989 | -0.43419 | dimension
Micro-Vickers | log_ratio | 54 | -0.0012486 | 0.005348 | -0.011731 | 0.0092335 | 54 | 54 | -0.0037886 | 0.00070702 | 0.0031807 | 0.0073839 | -0.017626 | -0.0058736 | 0.0059278 | 0.011962 | 0.00039789 | 0.76662 | -0.43419 | dimension
Vickers | absolute | 60 | -0.025958 | 0.36702 | -0.74532 | 0.6934 | 60 | 60 | -0.14365 | 0.08526 | 0.3023 | 0.41635 | -0.90045 | -0.56165 | 0.515 | 0.81734 | 0.00043301 | 0.01801 | -0.10597 | dimension
Vickers | relative_percent | 60 | -0.029036 | 0.22267 | -0.46547 | 0.4074 | 60 | 60 | -0.09246 | 0.029948 | 0.11203 | 0.31133 | -0.68935 | -0.22352 | 0.20175 | 0.55346 | 0.00013677 | 0.22638 | -0.57899 | dimension
Vickers | log_ratio | 60 | -0.00029036 | 0.0022267 | -0.0046548 | 0.004074 | 60 | 60 | -0.00092461 | 0.00029948 | 0.0011203 | 0.0031133 | -0.0068935 | -0.0022352 | 0.0020175 | 0.0055346 | 0.00052329 | 0.18653 | -0.57899 | dimension
Brinell | absolute | 36 | -4.7447 | 7.5532 | -19.549 | 10.0596 | 36 | 36 | -8.5264 | -2.1297 | 1.9087 | 11.7701 | -31.1767 | -5.972 | 1.3986 | 15.0792 | -0.011893 | 0.15359 | 0.1502 | hardness
Brinell | relative_percent | 36 | -1.602 | 1.8532 | -5.2342 | 2.0303 | 36 | 36 | -2.5276 | -0.87854 | 0.70572 | 2.7325 | -7.6785 | -2.2783 | 0.41933 | 3.1505 | 0.00037587 | 0.8563 | -0.30517 | hardness
Brinell | log_ratio | 36 | -0.016023 | 0.018544 | -0.052368 | 0.020323 | 36 | 36 | -0.025282 | -0.0087857 | 0.0070575 | 0.027346 | -0.076834 | -0.022784 | 0.0041951 | 0.031538 | 0.00054583 | 0.92965 | -0.30605 | hardness
Micro-Vickers | absolute | 54 | 2.7644 | 8.2198 | -13.3465 | 18.8753 | 54 | 54 | 0.011011 | 6.957 | 3.7488 | 12.4563 | -18.4567 | -6.3998 | 7.6352 | 30.3474 | 0.012114 | 0.00041051 | 0.73852 | hardness
Micro-Vickers | relative_percent | 54 | 0.21471 | 1.1264 | -1.993 | 2.4225 | 54 | 54 | -0.18748 | 0.7052 | 0.70445 | 1.5201 | -2.5715 | -1.2876 | 1.2738 | 3.5674 | 0.0018322 | 7.5771e-05 | 0.32159 | hardness
Micro-Vickers | log_ratio | 54 | 0.0021473 | 0.011265 | -0.019932 | 0.024226 | 54 | 54 | -0.0018748 | 0.0070525 | 0.0070446 | 0.015202 | -0.025717 | -0.012876 | 0.012739 | 0.035677 | 0.0059004 | 0.00023496 | 0.32159 | hardness
Vickers | absolute | 60 | 0.63642 | 3.2833 | -5.7989 | 7.0718 | 60 | 60 | -0.15204 | 1.5372 | 1.2079 | 4.8015 | -8.2562 | -2.1263 | 2.5391 | 10.795 | 0.0042249 | 0.0065021 | 0.70682 | hardness
Vickers | relative_percent | 60 | 0.040415 | 0.4353 | -0.81278 | 0.89361 | 60 | 60 | -0.071653 | 0.16731 | 0.20015 | 0.61652 | -1.108 | -0.37995 | 0.39002 | 1.3542 | 0.00059469 | 0.0037108 | 0.42115 | hardness
Vickers | log_ratio | 60 | 0.00040417 | 0.0043531 | -0.0081279 | 0.0089362 | 60 | 60 | -0.00071652 | 0.0016731 | 0.0020015 | 0.0061653 | -0.01108 | -0.0037995 | 0.0039002 | 0.013542 | 0.0022161 | 0.013623 | 0.42047 | hardness

## Bland_Altman_Points
pair_mean | difference | manual | automatic | method | representation | source_index | series_id | magnification | hardness_scale | hardness_level | applied_force_n | ball_diameter_mm | serial_number
211.6415 | 1.385 | 210.949 | 212.334 | Brinell | absolute | 114 | Brinell¦SN=293-250¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | 20X | HBW 1/5 | 1 | 49.05 | 1 | 293-250
211.7795 | 0.931 | 211.314 | 212.245 | Brinell | absolute | 115 | Brinell¦SN=293-250¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | 20X | HBW 1/5 | 1 | 49.05 | 1 | 293-250
210.097 | 1.945 | 209.1245 | 211.0695 | Brinell | absolute | 116 | Brinell¦SN=293-250¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | 20X | HBW 1/5 | 1 | 49.05 | 1 | 293-250
155.0702 | 0.8945 | 154.623 | 155.5175 | Brinell | absolute | 117 | Brinell¦SN=294-367¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | 20X | HBW 1/5 | 1 | 49.05 | 1 | 294-367
155.013 | 0.537 | 154.7445 | 155.2815 | Brinell | absolute | 118 | Brinell¦SN=294-367¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | 20X | HBW 1/5 | 1 | 49.05 | 1 | 294-367
154.569 | 0.378 | 154.38 | 154.758 | Brinell | absolute | 119 | Brinell¦SN=294-367¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | 20X | HBW 1/5 | 1 | 49.05 | 1 | 294-367
116.667 | 1.46 | 115.937 | 117.397 | Brinell | absolute | 120 | Brinell¦SN=290-736¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | 20X | HBW 1/5 | 1 | 49.05 | 1 | 290-736
119.6208 | 0.7955 | 119.223 | 120.0185 | Brinell | absolute | 121 | Brinell¦SN=290-736¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | 20X | HBW 1/5 | 1 | 49.05 | 1 | 290-736
116.0905 | 0.548 | 115.8165 | 116.3645 | Brinell | absolute | 122 | Brinell¦SN=290-736¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | 20X | HBW 1/5 | 1 | 49.05 | 1 | 290-736
466.3827 | 2.6145 | 465.0755 | 467.69 | Brinell | absolute | 123 | Brinell¦SN=293-250¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | 20X | HBW 1/30 | 1 | 294.3 | 1 | 293-250
465.523 | 1.338 | 464.854 | 466.192 | Brinell | absolute | 124 | Brinell¦SN=293-250¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | 20X | HBW 1/30 | 1 | 294.3 | 1 | 293-250
465.1343 | 1.6875 | 464.2905 | 465.978 | Brinell | absolute | 125 | Brinell¦SN=293-250¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | 20X | HBW 1/30 | 1 | 294.3 | 1 | 293-250
352.2265 | 0.6 | 351.9265 | 352.5265 | Brinell | absolute | 126 | Brinell¦SN=294-367¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | 20X | HBW 1/30 | 1 | 294.3 | 1 | 294-367
350.464 | 0.169 | 350.3795 | 350.5485 | Brinell | absolute | 127 | Brinell¦SN=294-367¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | 20X | HBW 1/30 | 1 | 294.3 | 1 | 294-367
349.9715 | 0.568 | 349.6875 | 350.2555 | Brinell | absolute | 128 | Brinell¦SN=294-367¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | 20X | HBW 1/30 | 1 | 294.3 | 1 | 294-367
255.632 | 1.79 | 254.737 | 256.527 | Brinell | absolute | 129 | Brinell¦SN=290-736¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | 20X | HBW 1/30 | 1 | 294.3 | 1 | 290-736
254.8057 | 0.1295 | 254.741 | 254.8705 | Brinell | absolute | 130 | Brinell¦SN=290-736¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | 20X | HBW 1/30 | 1 | 294.3 | 1 | 290-736
255.0825 | -0.528 | 255.3465 | 254.8185 | Brinell | absolute | 131 | Brinell¦SN=290-736¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | 20X | HBW 1/30 | 1 | 294.3 | 1 | 290-736
2219.7967 | 36.7695 | 2201.412 | 2238.1815 | Brinell | absolute | 132 | Brinell¦SN=812-006¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=20X | 20X | HBW 10/500 | 10 | 4905 | 10 | 812-006
2218.1002 | 19.7305 | 2208.235 | 2227.9655 | Brinell | absolute | 133 | Brinell¦SN=812-006¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=20X | 20X | HBW 10/500 | 10 | 4905 | 10 | 812-006
2218.2655 | 37.023 | 2199.754 | 2236.777 | Brinell | absolute | 134 | Brinell¦SN=812-006¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=20X | 20X | HBW 10/500 | 10 | 4905 | 10 | 812-006
1699.9185 | 25.788 | 1687.0245 | 1712.8125 | Brinell | absolute | 135 | Brinell¦SN=810-202¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=20X | 20X | HBW 10/500 | 10 | 4905 | 10 | 810-202
1691.926 | 14.126 | 1684.863 | 1698.989 | Brinell | absolute | 136 | Brinell¦SN=810-202¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=20X | 20X | HBW 10/500 | 10 | 4905 | 10 | 810-202
1703.2423 | 33.0245 | 1686.73 | 1719.7545 | Brinell | absolute | 137 | Brinell¦SN=810-202¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=20X | 20X | HBW 10/500 | 10 | 4905 | 10 | 810-202
1218.1195 | 25.655 | 1205.292 | 1230.947 | Brinell | absolute | 138 | Brinell¦SN=8301-50DW¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=10X | 10X | HBW 10/500 | 10 | 4905 | 10 | 8301-50DW
1233.5887 | 61.4805 | 1202.8485 | 1264.329 | Brinell | absolute | 139 | Brinell¦SN=8301-50DW¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=10X | 10X | HBW 10/500 | 10 | 4905 | 10 | 8301-50DW
1214.718 | 19.474 | 1204.981 | 1224.455 | Brinell | absolute | 140 | Brinell¦SN=8301-50DW¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=10X | 10X | HBW 10/500 | 10 | 4905 | 10 | 8301-50DW
4842.1218 | 19.8445 | 4832.1995 | 4852.044 | Brinell | absolute | 141 | Brinell¦SN=812-006¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | 20X | HBW 10/3000 | 10 | 29430 | 10 | 812-006
4865.6052 | 29.7465 | 4850.732 | 4880.4785 | Brinell | absolute | 142 | Brinell¦SN=812-006¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | 20X | HBW 10/3000 | 10 | 29430 | 10 | 812-006
4845.0905 | 11.32 | 4839.4305 | 4850.7505 | Brinell | absolute | 143 | Brinell¦SN=812-006¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | 20X | HBW 10/3000 | 10 | 29430 | 10 | 812-006
3813.5033 | 14.2695 | 3806.3685 | 3820.638 | Brinell | absolute | 144 | Brinell¦SN=810-202¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | 20X | HBW 10/3000 | 10 | 29430 | 10 | 810-202
3821.081 | 39.229 | 3801.4665 | 3840.6955 | Brinell | absolute | 145 | Brinell¦SN=810-202¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | 20X | HBW 10/3000 | 10 | 29430 | 10 | 810-202
3808.3908 | 2.9505 | 3806.9155 | 3809.866 | Brinell | absolute | 146 | Brinell¦SN=810-202¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | 20X | HBW 10/3000 | 10 | 29430 | 10 | 810-202
2659.7875 | 6.425 | 2656.575 | 2663 | Brinell | absolute | 147 | Brinell¦SN=8301-50DW¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | 20X | HBW 10/3000 | 10 | 29430 | 10 | 8301-50DW
2651.1435 | 11.694 | 2645.2965 | 2656.9905 | Brinell | absolute | 148 | Brinell¦SN=8301-50DW¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | 20X | HBW 10/3000 | 10 | 29430 | 10 | 8301-50DW
2646.8063 | 2.3575 | 2645.6275 | 2647.985 | Brinell | absolute | 149 | Brinell¦SN=8301-50DW¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | 20X | HBW 10/3000 | 10 | 29430 | 10 | 8301-50DW
211.6415 | 0.65441 | 210.949 | 212.334 | Brinell | relative_percent | 114 | Brinell¦SN=293-250¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | 20X | HBW 1/5 | 1 | 49.05 | 1 | 293-250
211.7795 | 0.43961 | 211.314 | 212.245 | Brinell | relative_percent | 115 | Brinell¦SN=293-250¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | 20X | HBW 1/5 | 1 | 49.05 | 1 | 293-250
210.097 | 0.92576 | 209.1245 | 211.0695 | Brinell | relative_percent | 116 | Brinell¦SN=293-250¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | 20X | HBW 1/5 | 1 | 49.05 | 1 | 293-250
155.0702 | 0.57684 | 154.623 | 155.5175 | Brinell | relative_percent | 117 | Brinell¦SN=294-367¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | 20X | HBW 1/5 | 1 | 49.05 | 1 | 294-367
155.013 | 0.34642 | 154.7445 | 155.2815 | Brinell | relative_percent | 118 | Brinell¦SN=294-367¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | 20X | HBW 1/5 | 1 | 49.05 | 1 | 294-367
154.569 | 0.24455 | 154.38 | 154.758 | Brinell | relative_percent | 119 | Brinell¦SN=294-367¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | 20X | HBW 1/5 | 1 | 49.05 | 1 | 294-367
116.667 | 1.2514 | 115.937 | 117.397 | Brinell | relative_percent | 120 | Brinell¦SN=290-736¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | 20X | HBW 1/5 | 1 | 49.05 | 1 | 290-736
119.6208 | 0.66502 | 119.223 | 120.0185 | Brinell | relative_percent | 121 | Brinell¦SN=290-736¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | 20X | HBW 1/5 | 1 | 49.05 | 1 | 290-736
116.0905 | 0.47205 | 115.8165 | 116.3645 | Brinell | relative_percent | 122 | Brinell¦SN=290-736¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | 20X | HBW 1/5 | 1 | 49.05 | 1 | 290-736
466.3827 | 0.56059 | 465.0755 | 467.69 | Brinell | relative_percent | 123 | Brinell¦SN=293-250¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | 20X | HBW 1/30 | 1 | 294.3 | 1 | 293-250
465.523 | 0.28742 | 464.854 | 466.192 | Brinell | relative_percent | 124 | Brinell¦SN=293-250¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | 20X | HBW 1/30 | 1 | 294.3 | 1 | 293-250
465.1343 | 0.3628 | 464.2905 | 465.978 | Brinell | relative_percent | 125 | Brinell¦SN=293-250¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | 20X | HBW 1/30 | 1 | 294.3 | 1 | 293-250
352.2265 | 0.17034 | 351.9265 | 352.5265 | Brinell | relative_percent | 126 | Brinell¦SN=294-367¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | 20X | HBW 1/30 | 1 | 294.3 | 1 | 294-367
350.464 | 0.048222 | 350.3795 | 350.5485 | Brinell | relative_percent | 127 | Brinell¦SN=294-367¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | 20X | HBW 1/30 | 1 | 294.3 | 1 | 294-367
349.9715 | 0.1623 | 349.6875 | 350.2555 | Brinell | relative_percent | 128 | Brinell¦SN=294-367¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | 20X | HBW 1/30 | 1 | 294.3 | 1 | 294-367
255.632 | 0.70023 | 254.737 | 256.527 | Brinell | relative_percent | 129 | Brinell¦SN=290-736¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | 20X | HBW 1/30 | 1 | 294.3 | 1 | 290-736
254.8057 | 0.050823 | 254.741 | 254.8705 | Brinell | relative_percent | 130 | Brinell¦SN=290-736¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | 20X | HBW 1/30 | 1 | 294.3 | 1 | 290-736
255.0825 | -0.20699 | 255.3465 | 254.8185 | Brinell | relative_percent | 131 | Brinell¦SN=290-736¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | 20X | HBW 1/30 | 1 | 294.3 | 1 | 290-736
2219.7967 | 1.6564 | 2201.412 | 2238.1815 | Brinell | relative_percent | 132 | Brinell¦SN=812-006¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=20X | 20X | HBW 10/500 | 10 | 4905 | 10 | 812-006
2218.1002 | 0.88952 | 2208.235 | 2227.9655 | Brinell | relative_percent | 133 | Brinell¦SN=812-006¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=20X | 20X | HBW 10/500 | 10 | 4905 | 10 | 812-006
2218.2655 | 1.669 | 2199.754 | 2236.777 | Brinell | relative_percent | 134 | Brinell¦SN=812-006¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=20X | 20X | HBW 10/500 | 10 | 4905 | 10 | 812-006
1699.9185 | 1.517 | 1687.0245 | 1712.8125 | Brinell | relative_percent | 135 | Brinell¦SN=810-202¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=20X | 20X | HBW 10/500 | 10 | 4905 | 10 | 810-202
1691.926 | 0.83491 | 1684.863 | 1698.989 | Brinell | relative_percent | 136 | Brinell¦SN=810-202¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=20X | 20X | HBW 10/500 | 10 | 4905 | 10 | 810-202
1703.2423 | 1.9389 | 1686.73 | 1719.7545 | Brinell | relative_percent | 137 | Brinell¦SN=810-202¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=20X | 20X | HBW 10/500 | 10 | 4905 | 10 | 810-202
1218.1195 | 2.1061 | 1205.292 | 1230.947 | Brinell | relative_percent | 138 | Brinell¦SN=8301-50DW¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=10X | 10X | HBW 10/500 | 10 | 4905 | 10 | 8301-50DW
1233.5887 | 4.9839 | 1202.8485 | 1264.329 | Brinell | relative_percent | 139 | Brinell¦SN=8301-50DW¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=10X | 10X | HBW 10/500 | 10 | 4905 | 10 | 8301-50DW
1214.718 | 1.6032 | 1204.981 | 1224.455 | Brinell | relative_percent | 140 | Brinell¦SN=8301-50DW¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=10X | 10X | HBW 10/500 | 10 | 4905 | 10 | 8301-50DW
4842.1218 | 0.40983 | 4832.1995 | 4852.044 | Brinell | relative_percent | 141 | Brinell¦SN=812-006¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | 20X | HBW 10/3000 | 10 | 29430 | 10 | 812-006
4865.6052 | 0.61136 | 4850.732 | 4880.4785 | Brinell | relative_percent | 142 | Brinell¦SN=812-006¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | 20X | HBW 10/3000 | 10 | 29430 | 10 | 812-006
4845.0905 | 0.23364 | 4839.4305 | 4850.7505 | Brinell | relative_percent | 143 | Brinell¦SN=812-006¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | 20X | HBW 10/3000 | 10 | 29430 | 10 | 812-006
3813.5033 | 0.37418 | 3806.3685 | 3820.638 | Brinell | relative_percent | 144 | Brinell¦SN=810-202¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | 20X | HBW 10/3000 | 10 | 29430 | 10 | 810-202
3821.081 | 1.0266 | 3801.4665 | 3840.6955 | Brinell | relative_percent | 145 | Brinell¦SN=810-202¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | 20X | HBW 10/3000 | 10 | 29430 | 10 | 810-202
3808.3908 | 0.077474 | 3806.9155 | 3809.866 | Brinell | relative_percent | 146 | Brinell¦SN=810-202¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | 20X | HBW 10/3000 | 10 | 29430 | 10 | 810-202
2659.7875 | 0.24156 | 2656.575 | 2663 | Brinell | relative_percent | 147 | Brinell¦SN=8301-50DW¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | 20X | HBW 10/3000 | 10 | 29430 | 10 | 8301-50DW
2651.1435 | 0.44109 | 2645.2965 | 2656.9905 | Brinell | relative_percent | 148 | Brinell¦SN=8301-50DW¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | 20X | HBW 10/3000 | 10 | 29430 | 10 | 8301-50DW
2646.8063 | 0.08907 | 2645.6275 | 2647.985 | Brinell | relative_percent | 149 | Brinell¦SN=8301-50DW¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | 20X | HBW 10/3000 | 10 | 29430 | 10 | 8301-50DW
5.3549 | 0.0065441 | 210.949 | 212.334 | Brinell | log_ratio | 114 | Brinell¦SN=293-250¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | 20X | HBW 1/5 | 1 | 49.05 | 1 | 293-250
5.3555 | 0.0043961 | 211.314 | 212.245 | Brinell | log_ratio | 115 | Brinell¦SN=293-250¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | 20X | HBW 1/5 | 1 | 49.05 | 1 | 293-250
5.3476 | 0.0092577 | 209.1245 | 211.0695 | Brinell | log_ratio | 116 | Brinell¦SN=293-250¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | 20X | HBW 1/5 | 1 | 49.05 | 1 | 293-250
5.0439 | 0.0057684 | 154.623 | 155.5175 | Brinell | log_ratio | 117 | Brinell¦SN=294-367¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | 20X | HBW 1/5 | 1 | 49.05 | 1 | 294-367
5.0435 | 0.0034642 | 154.7445 | 155.2815 | Brinell | log_ratio | 118 | Brinell¦SN=294-367¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | 20X | HBW 1/5 | 1 | 49.05 | 1 | 294-367
5.0406 | 0.0024455 | 154.38 | 154.758 | Brinell | log_ratio | 119 | Brinell¦SN=294-367¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | 20X | HBW 1/5 | 1 | 49.05 | 1 | 294-367
4.7593 | 0.012514 | 115.937 | 117.397 | Brinell | log_ratio | 120 | Brinell¦SN=290-736¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | 20X | HBW 1/5 | 1 | 49.05 | 1 | 290-736
4.7843 | 0.0066502 | 119.223 | 120.0185 | Brinell | log_ratio | 121 | Brinell¦SN=290-736¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | 20X | HBW 1/5 | 1 | 49.05 | 1 | 290-736
4.7544 | 0.0047205 | 115.8165 | 116.3645 | Brinell | log_ratio | 122 | Brinell¦SN=290-736¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | 20X | HBW 1/5 | 1 | 49.05 | 1 | 290-736
6.145 | 0.0056059 | 465.0755 | 467.69 | Brinell | log_ratio | 123 | Brinell¦SN=293-250¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | 20X | HBW 1/30 | 1 | 294.3 | 1 | 293-250
6.1432 | 0.0028742 | 464.854 | 466.192 | Brinell | log_ratio | 124 | Brinell¦SN=293-250¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | 20X | HBW 1/30 | 1 | 294.3 | 1 | 293-250
6.1423 | 0.003628 | 464.2905 | 465.978 | Brinell | log_ratio | 125 | Brinell¦SN=293-250¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | 20X | HBW 1/30 | 1 | 294.3 | 1 | 293-250
5.8643 | 0.0017034 | 351.9265 | 352.5265 | Brinell | log_ratio | 126 | Brinell¦SN=294-367¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | 20X | HBW 1/30 | 1 | 294.3 | 1 | 294-367
5.8593 | 0.00048222 | 350.3795 | 350.5485 | Brinell | log_ratio | 127 | Brinell¦SN=294-367¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | 20X | HBW 1/30 | 1 | 294.3 | 1 | 294-367
5.8579 | 0.001623 | 349.6875 | 350.2555 | Brinell | log_ratio | 128 | Brinell¦SN=294-367¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | 20X | HBW 1/30 | 1 | 294.3 | 1 | 294-367
5.5437 | 0.0070023 | 254.737 | 256.527 | Brinell | log_ratio | 129 | Brinell¦SN=290-736¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | 20X | HBW 1/30 | 1 | 294.3 | 1 | 290-736
5.5405 | 0.00050823 | 254.741 | 254.8705 | Brinell | log_ratio | 130 | Brinell¦SN=290-736¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | 20X | HBW 1/30 | 1 | 294.3 | 1 | 290-736
5.5416 | -0.0020699 | 255.3465 | 254.8185 | Brinell | log_ratio | 131 | Brinell¦SN=290-736¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | 20X | HBW 1/30 | 1 | 294.3 | 1 | 290-736
7.7051 | 0.016565 | 2201.412 | 2238.1815 | Brinell | log_ratio | 132 | Brinell¦SN=812-006¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=20X | 20X | HBW 10/500 | 10 | 4905 | 10 | 812-006
7.7044 | 0.0088953 | 2208.235 | 2227.9655 | Brinell | log_ratio | 133 | Brinell¦SN=812-006¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=20X | 20X | HBW 10/500 | 10 | 4905 | 10 | 812-006
7.7044 | 0.01669 | 2199.754 | 2236.777 | Brinell | log_ratio | 134 | Brinell¦SN=812-006¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=20X | 20X | HBW 10/500 | 10 | 4905 | 10 | 812-006
7.4383 | 0.01517 | 1687.0245 | 1712.8125 | Brinell | log_ratio | 135 | Brinell¦SN=810-202¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=20X | 20X | HBW 10/500 | 10 | 4905 | 10 | 810-202
7.4336 | 0.0083491 | 1684.863 | 1698.989 | Brinell | log_ratio | 136 | Brinell¦SN=810-202¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=20X | 20X | HBW 10/500 | 10 | 4905 | 10 | 810-202
7.4402 | 0.01939 | 1686.73 | 1719.7545 | Brinell | log_ratio | 137 | Brinell¦SN=810-202¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=20X | 20X | HBW 10/500 | 10 | 4905 | 10 | 810-202
7.105 | 0.021062 | 1205.292 | 1230.947 | Brinell | log_ratio | 138 | Brinell¦SN=8301-50DW¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=10X | 10X | HBW 10/500 | 10 | 4905 | 10 | 8301-50DW
7.1174 | 0.049849 | 1202.8485 | 1264.329 | Brinell | log_ratio | 139 | Brinell¦SN=8301-50DW¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=10X | 10X | HBW 10/500 | 10 | 4905 | 10 | 8301-50DW
7.1022 | 0.016032 | 1204.981 | 1224.455 | Brinell | log_ratio | 140 | Brinell¦SN=8301-50DW¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=10X | 10X | HBW 10/500 | 10 | 4905 | 10 | 8301-50DW
8.4851 | 0.0040983 | 4832.1995 | 4852.044 | Brinell | log_ratio | 141 | Brinell¦SN=812-006¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | 20X | HBW 10/3000 | 10 | 29430 | 10 | 812-006
8.4899 | 0.0061136 | 4850.732 | 4880.4785 | Brinell | log_ratio | 142 | Brinell¦SN=812-006¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | 20X | HBW 10/3000 | 10 | 29430 | 10 | 812-006
8.4857 | 0.0023364 | 4839.4305 | 4850.7505 | Brinell | log_ratio | 143 | Brinell¦SN=812-006¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | 20X | HBW 10/3000 | 10 | 29430 | 10 | 812-006
8.2463 | 0.0037418 | 3806.3685 | 3820.638 | Brinell | log_ratio | 144 | Brinell¦SN=810-202¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | 20X | HBW 10/3000 | 10 | 29430 | 10 | 810-202
8.2483 | 0.010267 | 3801.4665 | 3840.6955 | Brinell | log_ratio | 145 | Brinell¦SN=810-202¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | 20X | HBW 10/3000 | 10 | 29430 | 10 | 810-202
8.245 | 0.00077474 | 3806.9155 | 3809.866 | Brinell | log_ratio | 146 | Brinell¦SN=810-202¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | 20X | HBW 10/3000 | 10 | 29430 | 10 | 810-202
7.886 | 0.0024156 | 2656.575 | 2663 | Brinell | log_ratio | 147 | Brinell¦SN=8301-50DW¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | 20X | HBW 10/3000 | 10 | 29430 | 10 | 8301-50DW
7.8827 | 0.0044109 | 2645.2965 | 2656.9905 | Brinell | log_ratio | 148 | Brinell¦SN=8301-50DW¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | 20X | HBW 10/3000 | 10 | 29430 | 10 | 8301-50DW
7.8811 | 0.0008907 | 2645.6275 | 2647.985 | Brinell | log_ratio | 149 | Brinell¦SN=8301-50DW¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | 20X | HBW 10/3000 | 10 | 29430 | 10 | 8301-50DW
44.2965 | 0.163 | 44.215 | 44.378 | Micro-Vickers | absolute | 60 | Micro-Vickers¦SN=710-404¦scale=HV 0.05¦F=0.4905N¦mag=50X | 50X | HV 0.05 | 0.05 | 0.4905 |  | 710-404
43.6457 | 0.0865 | 43.6025 | 43.689 | Micro-Vickers | absolute | 61 | Micro-Vickers¦SN=710-404¦scale=HV 0.05¦F=0.4905N¦mag=50X | 50X | HV 0.05 | 0.05 | 0.4905 |  | 710-404
42.27 | 0.16 | 42.19 | 42.35 | Micro-Vickers | absolute | 62 | Micro-Vickers¦SN=710-404¦scale=HV 0.05¦F=0.4905N¦mag=50X | 50X | HV 0.05 | 0.05 | 0.4905 |  | 710-404
28.6085 | 0.032 | 28.5925 | 28.6245 | Micro-Vickers | absolute | 63 | Micro-Vickers¦SN=711-081¦scale=HV 0.05¦F=0.4905N¦mag=50X | 50X | HV 0.05 | 0.05 | 0.4905 |  | 711-081
29.1285 | 0.092 | 29.0825 | 29.1745 | Micro-Vickers | absolute | 64 | Micro-Vickers¦SN=711-081¦scale=HV 0.05¦F=0.4905N¦mag=50X | 50X | HV 0.05 | 0.05 | 0.4905 |  | 711-081
28.6552 | 0.0755 | 28.6175 | 28.693 | Micro-Vickers | absolute | 65 | Micro-Vickers¦SN=711-081¦scale=HV 0.05¦F=0.4905N¦mag=50X | 50X | HV 0.05 | 0.05 | 0.4905 |  | 711-081
20.873 | 0.062 | 20.842 | 20.904 | Micro-Vickers | absolute | 66 | Micro-Vickers¦SN=710-655¦scale=HV 0.05¦F=0.4905N¦mag=50X | 50X | HV 0.05 | 0.05 | 0.4905 |  | 710-655
20.7473 | -0.0305 | 20.7625 | 20.732 | Micro-Vickers | absolute | 67 | Micro-Vickers¦SN=710-655¦scale=HV 0.05¦F=0.4905N¦mag=50X | 50X | HV 0.05 | 0.05 | 0.4905 |  | 710-655
20.982 | 0.078 | 20.943 | 21.021 | Micro-Vickers | absolute | 68 | Micro-Vickers¦SN=710-655¦scale=HV 0.05¦F=0.4905N¦mag=50X | 50X | HV 0.05 | 0.05 | 0.4905 |  | 710-655
94.4515 | 0.153 | 94.375 | 94.528 | Micro-Vickers | absolute | 69 | Micro-Vickers¦SN=711-081¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 711-081
93.2335 | 0.24 | 93.1135 | 93.3535 | Micro-Vickers | absolute | 70 | Micro-Vickers¦SN=711-081¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 711-081
94.2258 | 0.0385 | 94.2065 | 94.245 | Micro-Vickers | absolute | 71 | Micro-Vickers¦SN=711-081¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 711-081
53.9552 | -0.1145 | 54.0125 | 53.898 | Micro-Vickers | absolute | 72 | Micro-Vickers¦SN=706-437¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 706-437
53.7925 | 0.107 | 53.739 | 53.846 | Micro-Vickers | absolute | 73 | Micro-Vickers¦SN=706-437¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 706-437
54.1895 | 0.141 | 54.119 | 54.26 | Micro-Vickers | absolute | 74 | Micro-Vickers¦SN=706-437¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 706-437
42.971 | 0.162 | 42.89 | 43.052 | Micro-Vickers | absolute | 75 | Micro-Vickers¦SN=711-246¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 711-246
42.3867 | -0.0745 | 42.424 | 42.3495 | Micro-Vickers | absolute | 76 | Micro-Vickers¦SN=711-246¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 711-246
42.7383 | 0.4865 | 42.495 | 42.9815 | Micro-Vickers | absolute | 77 | Micro-Vickers¦SN=711-246¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 711-246
35.0763 | -0.2765 | 35.2145 | 34.938 | Micro-Vickers | absolute | 78 | Micro-Vickers¦SN=710-106¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 710-106
35.2705 | 0.091 | 35.225 | 35.316 | Micro-Vickers | absolute | 79 | Micro-Vickers¦SN=710-106¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 710-106
34.8327 | -0.1535 | 34.9095 | 34.756 | Micro-Vickers | absolute | 80 | Micro-Vickers¦SN=710-106¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 710-106
30.867 | 0.129 | 30.8025 | 30.9315 | Micro-Vickers | absolute | 81 | Micro-Vickers¦SN=709-293¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 709-293
30.7948 | 0.0545 | 30.7675 | 30.822 | Micro-Vickers | absolute | 82 | Micro-Vickers¦SN=709-293¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 709-293
30.5547 | -0.1145 | 30.612 | 30.4975 | Micro-Vickers | absolute | 83 | Micro-Vickers¦SN=709-293¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 709-293
134.053 | 0.27 | 133.918 | 134.188 | Micro-Vickers | absolute | 84 | Micro-Vickers¦SN=711-081¦scale=HV 1¦F=9.81N¦mag=20X | 20X | HV 1 | 1 | 9.81 |  | 711-081
134.2715 | 0.109 | 134.217 | 134.326 | Micro-Vickers | absolute | 85 | Micro-Vickers¦SN=711-081¦scale=HV 1¦F=9.81N¦mag=20X | 20X | HV 1 | 1 | 9.81 |  | 711-081
133.4563 | 0.0035 | 133.4545 | 133.458 | Micro-Vickers | absolute | 86 | Micro-Vickers¦SN=711-081¦scale=HV 1¦F=9.81N¦mag=20X | 20X | HV 1 | 1 | 9.81 |  | 711-081
77.623 | -0.213 | 77.7295 | 77.5165 | Micro-Vickers | absolute | 87 | Micro-Vickers¦SN=706-437¦scale=HV 1¦F=9.81N¦mag=50X | 50X | HV 1 | 1 | 9.81 |  | 706-437
75.4437 | -0.4495 | 75.6685 | 75.219 | Micro-Vickers | absolute | 88 | Micro-Vickers¦SN=706-437¦scale=HV 1¦F=9.81N¦mag=50X | 50X | HV 1 | 1 | 9.81 |  | 706-437
76.419 | -0.422 | 76.63 | 76.208 | Micro-Vickers | absolute | 89 | Micro-Vickers¦SN=706-437¦scale=HV 1¦F=9.81N¦mag=50X | 50X | HV 1 | 1 | 9.81 |  | 706-437
60.9123 | 0.0205 | 60.902 | 60.9225 | Micro-Vickers | absolute | 90 | Micro-Vickers¦SN=711-246¦scale=HV 1¦F=9.81N¦mag=50X | 50X | HV 1 | 1 | 9.81 |  | 711-246
60.2545 | -0.63 | 60.5695 | 59.9395 | Micro-Vickers | absolute | 91 | Micro-Vickers¦SN=711-246¦scale=HV 1¦F=9.81N¦mag=50X | 50X | HV 1 | 1 | 9.81 |  | 711-246
60.0258 | -0.1135 | 60.0825 | 59.969 | Micro-Vickers | absolute | 92 | Micro-Vickers¦SN=711-246¦scale=HV 1¦F=9.81N¦mag=50X | 50X | HV 1 | 1 | 9.81 |  | 711-246
50.4868 | -0.3625 | 50.668 | 50.3055 | Micro-Vickers | absolute | 93 | Micro-Vickers¦SN=710-106¦scale=HV 1¦F=9.81N¦mag=50X | 50X | HV 1 | 1 | 9.81 |  | 710-106
50.197 | -0.279 | 50.3365 | 50.0575 | Micro-Vickers | absolute | 94 | Micro-Vickers¦SN=710-106¦scale=HV 1¦F=9.81N¦mag=50X | 50X | HV 1 | 1 | 9.81 |  | 710-106
49.879 | -0.331 | 50.0445 | 49.7135 | Micro-Vickers | absolute | 95 | Micro-Vickers¦SN=710-106¦scale=HV 1¦F=9.81N¦mag=50X | 50X | HV 1 | 1 | 9.81 |  | 710-106
44.696 | -0.81 | 45.101 | 44.291 | Micro-Vickers | absolute | 96 | Micro-Vickers¦SN=709-293¦scale=HV 1¦F=9.81N¦mag=50X | 50X | HV 1 | 1 | 9.81 |  | 709-293
44.5898 | -0.8435 | 45.0115 | 44.168 | Micro-Vickers | absolute | 97 | Micro-Vickers¦SN=709-293¦scale=HV 1¦F=9.81N¦mag=50X | 50X | HV 1 | 1 | 9.81 |  | 709-293
44.9592 | -0.5215 | 45.22 | 44.6985 | Micro-Vickers | absolute | 98 | Micro-Vickers¦SN=709-293¦scale=HV 1¦F=9.81N¦mag=50X | 50X | HV 1 | 1 | 9.81 |  | 709-293
190.0435 | 0.202 | 189.9425 | 190.1445 | Micro-Vickers | absolute | 99 | Micro-Vickers¦SN=711-081¦scale=HV 2¦F=19.62N¦mag=20X | 20X | HV 2 | 2 | 19.62 |  | 711-081
187.8092 | 0.4245 | 187.597 | 188.0215 | Micro-Vickers | absolute | 100 | Micro-Vickers¦SN=711-081¦scale=HV 2¦F=19.62N¦mag=20X | 20X | HV 2 | 2 | 19.62 |  | 711-081
190.6082 | 0.4015 | 190.4075 | 190.809 | Micro-Vickers | absolute | 101 | Micro-Vickers¦SN=711-081¦scale=HV 2¦F=19.62N¦mag=20X | 20X | HV 2 | 2 | 19.62 |  | 711-081
108.5952 | 0.0475 | 108.5715 | 108.619 | Micro-Vickers | absolute | 102 | Micro-Vickers¦SN=706-437¦scale=HV 2¦F=19.62N¦mag=20X | 20X | HV 2 | 2 | 19.62 |  | 706-437
107.93 | -0.06 | 107.96 | 107.9 | Micro-Vickers | absolute | 103 | Micro-Vickers¦SN=706-437¦scale=HV 2¦F=19.62N¦mag=20X | 20X | HV 2 | 2 | 19.62 |  | 706-437
107.6757 | 0.0915 | 107.63 | 107.7215 | Micro-Vickers | absolute | 104 | Micro-Vickers¦SN=706-437¦scale=HV 2¦F=19.62N¦mag=20X | 20X | HV 2 | 2 | 19.62 |  | 706-437
84.6175 | -0.109 | 84.672 | 84.563 | Micro-Vickers | absolute | 105 | Micro-Vickers¦SN=711-246¦scale=HV 2¦F=19.62N¦mag=50X | 50X | HV 2 | 2 | 19.62 |  | 711-246
85.024 | -0.06 | 85.054 | 84.994 | Micro-Vickers | absolute | 106 | Micro-Vickers¦SN=711-246¦scale=HV 2¦F=19.62N¦mag=50X | 50X | HV 2 | 2 | 19.62 |  | 711-246
85.0447 | -0.0185 | 85.054 | 85.0355 | Micro-Vickers | absolute | 107 | Micro-Vickers¦SN=711-246¦scale=HV 2¦F=19.62N¦mag=50X | 50X | HV 2 | 2 | 19.62 |  | 711-246
70.7807 | -0.0415 | 70.8015 | 70.76 | Micro-Vickers | absolute | 108 | Micro-Vickers¦SN=710-106¦scale=HV 2¦F=19.62N¦mag=50X | 50X | HV 2 | 2 | 19.62 |  | 710-106
70.7103 | -0.3465 | 70.8835 | 70.537 | Micro-Vickers | absolute | 109 | Micro-Vickers¦SN=710-106¦scale=HV 2¦F=19.62N¦mag=50X | 50X | HV 2 | 2 | 19.62 |  | 710-106
70.5553 | -0.3535 | 70.732 | 70.3785 | Micro-Vickers | absolute | 110 | Micro-Vickers¦SN=710-106¦scale=HV 2¦F=19.62N¦mag=50X | 50X | HV 2 | 2 | 19.62 |  | 710-106
63.7032 | 0.0875 | 63.6595 | 63.747 | Micro-Vickers | absolute | 111 | Micro-Vickers¦SN=709-293¦scale=HV 2¦F=19.62N¦mag=50X | 50X | HV 2 | 2 | 19.62 |  | 709-293
63.2482 | -0.0135 | 63.255 | 63.2415 | Micro-Vickers | absolute | 112 | Micro-Vickers¦SN=709-293¦scale=HV 2¦F=19.62N¦mag=50X | 50X | HV 2 | 2 | 19.62 |  | 709-293
62.6977 | -0.3245 | 62.86 | 62.5355 | Micro-Vickers | absolute | 113 | Micro-Vickers¦SN=709-293¦scale=HV 2¦F=19.62N¦mag=50X | 50X | HV 2 | 2 | 19.62 |  | 709-293
44.2965 | 0.36797 | 44.215 | 44.378 | Micro-Vickers | relative_percent | 60 | Micro-Vickers¦SN=710-404¦scale=HV 0.05¦F=0.4905N¦mag=50X | 50X | HV 0.05 | 0.05 | 0.4905 |  | 710-404
43.6457 | 0.19819 | 43.6025 | 43.689 | Micro-Vickers | relative_percent | 61 | Micro-Vickers¦SN=710-404¦scale=HV 0.05¦F=0.4905N¦mag=50X | 50X | HV 0.05 | 0.05 | 0.4905 |  | 710-404
42.27 | 0.37852 | 42.19 | 42.35 | Micro-Vickers | relative_percent | 62 | Micro-Vickers¦SN=710-404¦scale=HV 0.05¦F=0.4905N¦mag=50X | 50X | HV 0.05 | 0.05 | 0.4905 |  | 710-404
28.6085 | 0.11185 | 28.5925 | 28.6245 | Micro-Vickers | relative_percent | 63 | Micro-Vickers¦SN=711-081¦scale=HV 0.05¦F=0.4905N¦mag=50X | 50X | HV 0.05 | 0.05 | 0.4905 |  | 711-081
29.1285 | 0.31584 | 29.0825 | 29.1745 | Micro-Vickers | relative_percent | 64 | Micro-Vickers¦SN=711-081¦scale=HV 0.05¦F=0.4905N¦mag=50X | 50X | HV 0.05 | 0.05 | 0.4905 |  | 711-081
28.6552 | 0.26348 | 28.6175 | 28.693 | Micro-Vickers | relative_percent | 65 | Micro-Vickers¦SN=711-081¦scale=HV 0.05¦F=0.4905N¦mag=50X | 50X | HV 0.05 | 0.05 | 0.4905 |  | 711-081
20.873 | 0.29703 | 20.842 | 20.904 | Micro-Vickers | relative_percent | 66 | Micro-Vickers¦SN=710-655¦scale=HV 0.05¦F=0.4905N¦mag=50X | 50X | HV 0.05 | 0.05 | 0.4905 |  | 710-655
20.7473 | -0.14701 | 20.7625 | 20.732 | Micro-Vickers | relative_percent | 67 | Micro-Vickers¦SN=710-655¦scale=HV 0.05¦F=0.4905N¦mag=50X | 50X | HV 0.05 | 0.05 | 0.4905 |  | 710-655
20.982 | 0.37175 | 20.943 | 21.021 | Micro-Vickers | relative_percent | 68 | Micro-Vickers¦SN=710-655¦scale=HV 0.05¦F=0.4905N¦mag=50X | 50X | HV 0.05 | 0.05 | 0.4905 |  | 710-655
94.4515 | 0.16199 | 94.375 | 94.528 | Micro-Vickers | relative_percent | 69 | Micro-Vickers¦SN=711-081¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 711-081
93.2335 | 0.25742 | 93.1135 | 93.3535 | Micro-Vickers | relative_percent | 70 | Micro-Vickers¦SN=711-081¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 711-081
94.2258 | 0.040859 | 94.2065 | 94.245 | Micro-Vickers | relative_percent | 71 | Micro-Vickers¦SN=711-081¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 711-081
53.9552 | -0.21221 | 54.0125 | 53.898 | Micro-Vickers | relative_percent | 72 | Micro-Vickers¦SN=706-437¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 706-437
53.7925 | 0.19891 | 53.739 | 53.846 | Micro-Vickers | relative_percent | 73 | Micro-Vickers¦SN=706-437¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 706-437
54.1895 | 0.2602 | 54.119 | 54.26 | Micro-Vickers | relative_percent | 74 | Micro-Vickers¦SN=706-437¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 706-437
42.971 | 0.377 | 42.89 | 43.052 | Micro-Vickers | relative_percent | 75 | Micro-Vickers¦SN=711-246¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 711-246
42.3867 | -0.17576 | 42.424 | 42.3495 | Micro-Vickers | relative_percent | 76 | Micro-Vickers¦SN=711-246¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 711-246
42.7383 | 1.1383 | 42.495 | 42.9815 | Micro-Vickers | relative_percent | 77 | Micro-Vickers¦SN=711-246¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 711-246
35.0763 | -0.78828 | 35.2145 | 34.938 | Micro-Vickers | relative_percent | 78 | Micro-Vickers¦SN=710-106¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 710-106
35.2705 | 0.25801 | 35.225 | 35.316 | Micro-Vickers | relative_percent | 79 | Micro-Vickers¦SN=710-106¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 710-106
34.8327 | -0.44068 | 34.9095 | 34.756 | Micro-Vickers | relative_percent | 80 | Micro-Vickers¦SN=710-106¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 710-106
30.867 | 0.41792 | 30.8025 | 30.9315 | Micro-Vickers | relative_percent | 81 | Micro-Vickers¦SN=709-293¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 709-293
30.7948 | 0.17698 | 30.7675 | 30.822 | Micro-Vickers | relative_percent | 82 | Micro-Vickers¦SN=709-293¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 709-293
30.5547 | -0.37474 | 30.612 | 30.4975 | Micro-Vickers | relative_percent | 83 | Micro-Vickers¦SN=709-293¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 709-293
134.053 | 0.20141 | 133.918 | 134.188 | Micro-Vickers | relative_percent | 84 | Micro-Vickers¦SN=711-081¦scale=HV 1¦F=9.81N¦mag=20X | 20X | HV 1 | 1 | 9.81 |  | 711-081
134.2715 | 0.081179 | 134.217 | 134.326 | Micro-Vickers | relative_percent | 85 | Micro-Vickers¦SN=711-081¦scale=HV 1¦F=9.81N¦mag=20X | 20X | HV 1 | 1 | 9.81 |  | 711-081
133.4563 | 0.0026226 | 133.4545 | 133.458 | Micro-Vickers | relative_percent | 86 | Micro-Vickers¦SN=711-081¦scale=HV 1¦F=9.81N¦mag=20X | 20X | HV 1 | 1 | 9.81 |  | 711-081
77.623 | -0.2744 | 77.7295 | 77.5165 | Micro-Vickers | relative_percent | 87 | Micro-Vickers¦SN=706-437¦scale=HV 1¦F=9.81N¦mag=50X | 50X | HV 1 | 1 | 9.81 |  | 706-437
75.4437 | -0.59581 | 75.6685 | 75.219 | Micro-Vickers | relative_percent | 88 | Micro-Vickers¦SN=706-437¦scale=HV 1¦F=9.81N¦mag=50X | 50X | HV 1 | 1 | 9.81 |  | 706-437
76.419 | -0.55222 | 76.63 | 76.208 | Micro-Vickers | relative_percent | 89 | Micro-Vickers¦SN=706-437¦scale=HV 1¦F=9.81N¦mag=50X | 50X | HV 1 | 1 | 9.81 |  | 706-437
60.9123 | 0.033655 | 60.902 | 60.9225 | Micro-Vickers | relative_percent | 90 | Micro-Vickers¦SN=711-246¦scale=HV 1¦F=9.81N¦mag=50X | 50X | HV 1 | 1 | 9.81 |  | 711-246
60.2545 | -1.0456 | 60.5695 | 59.9395 | Micro-Vickers | relative_percent | 91 | Micro-Vickers¦SN=711-246¦scale=HV 1¦F=9.81N¦mag=50X | 50X | HV 1 | 1 | 9.81 |  | 711-246
60.0258 | -0.18909 | 60.0825 | 59.969 | Micro-Vickers | relative_percent | 92 | Micro-Vickers¦SN=711-246¦scale=HV 1¦F=9.81N¦mag=50X | 50X | HV 1 | 1 | 9.81 |  | 711-246
50.4868 | -0.71801 | 50.668 | 50.3055 | Micro-Vickers | relative_percent | 93 | Micro-Vickers¦SN=710-106¦scale=HV 1¦F=9.81N¦mag=50X | 50X | HV 1 | 1 | 9.81 |  | 710-106
50.197 | -0.55581 | 50.3365 | 50.0575 | Micro-Vickers | relative_percent | 94 | Micro-Vickers¦SN=710-106¦scale=HV 1¦F=9.81N¦mag=50X | 50X | HV 1 | 1 | 9.81 |  | 710-106
49.879 | -0.66361 | 50.0445 | 49.7135 | Micro-Vickers | relative_percent | 95 | Micro-Vickers¦SN=710-106¦scale=HV 1¦F=9.81N¦mag=50X | 50X | HV 1 | 1 | 9.81 |  | 710-106
44.696 | -1.8122 | 45.101 | 44.291 | Micro-Vickers | relative_percent | 96 | Micro-Vickers¦SN=709-293¦scale=HV 1¦F=9.81N¦mag=50X | 50X | HV 1 | 1 | 9.81 |  | 709-293
44.5898 | -1.8917 | 45.0115 | 44.168 | Micro-Vickers | relative_percent | 97 | Micro-Vickers¦SN=709-293¦scale=HV 1¦F=9.81N¦mag=50X | 50X | HV 1 | 1 | 9.81 |  | 709-293
44.9592 | -1.1599 | 45.22 | 44.6985 | Micro-Vickers | relative_percent | 98 | Micro-Vickers¦SN=709-293¦scale=HV 1¦F=9.81N¦mag=50X | 50X | HV 1 | 1 | 9.81 |  | 709-293
190.0435 | 0.10629 | 189.9425 | 190.1445 | Micro-Vickers | relative_percent | 99 | Micro-Vickers¦SN=711-081¦scale=HV 2¦F=19.62N¦mag=20X | 20X | HV 2 | 2 | 19.62 |  | 711-081
187.8092 | 0.22603 | 187.597 | 188.0215 | Micro-Vickers | relative_percent | 100 | Micro-Vickers¦SN=711-081¦scale=HV 2¦F=19.62N¦mag=20X | 20X | HV 2 | 2 | 19.62 |  | 711-081
190.6082 | 0.21064 | 190.4075 | 190.809 | Micro-Vickers | relative_percent | 101 | Micro-Vickers¦SN=711-081¦scale=HV 2¦F=19.62N¦mag=20X | 20X | HV 2 | 2 | 19.62 |  | 711-081
108.5952 | 0.04374 | 108.5715 | 108.619 | Micro-Vickers | relative_percent | 102 | Micro-Vickers¦SN=706-437¦scale=HV 2¦F=19.62N¦mag=20X | 20X | HV 2 | 2 | 19.62 |  | 706-437
107.93 | -0.055592 | 107.96 | 107.9 | Micro-Vickers | relative_percent | 103 | Micro-Vickers¦SN=706-437¦scale=HV 2¦F=19.62N¦mag=20X | 20X | HV 2 | 2 | 19.62 |  | 706-437
107.6757 | 0.084977 | 107.63 | 107.7215 | Micro-Vickers | relative_percent | 104 | Micro-Vickers¦SN=706-437¦scale=HV 2¦F=19.62N¦mag=20X | 20X | HV 2 | 2 | 19.62 |  | 706-437
84.6175 | -0.12881 | 84.672 | 84.563 | Micro-Vickers | relative_percent | 105 | Micro-Vickers¦SN=711-246¦scale=HV 2¦F=19.62N¦mag=50X | 50X | HV 2 | 2 | 19.62 |  | 711-246
85.024 | -0.070568 | 85.054 | 84.994 | Micro-Vickers | relative_percent | 106 | Micro-Vickers¦SN=711-246¦scale=HV 2¦F=19.62N¦mag=50X | 50X | HV 2 | 2 | 19.62 |  | 711-246
85.0447 | -0.021753 | 85.054 | 85.0355 | Micro-Vickers | relative_percent | 107 | Micro-Vickers¦SN=711-246¦scale=HV 2¦F=19.62N¦mag=50X | 50X | HV 2 | 2 | 19.62 |  | 711-246
70.7807 | -0.058632 | 70.8015 | 70.76 | Micro-Vickers | relative_percent | 108 | Micro-Vickers¦SN=710-106¦scale=HV 2¦F=19.62N¦mag=50X | 50X | HV 2 | 2 | 19.62 |  | 710-106
70.7103 | -0.49003 | 70.8835 | 70.537 | Micro-Vickers | relative_percent | 109 | Micro-Vickers¦SN=710-106¦scale=HV 2¦F=19.62N¦mag=50X | 50X | HV 2 | 2 | 19.62 |  | 710-106
70.5553 | -0.50103 | 70.732 | 70.3785 | Micro-Vickers | relative_percent | 110 | Micro-Vickers¦SN=710-106¦scale=HV 2¦F=19.62N¦mag=50X | 50X | HV 2 | 2 | 19.62 |  | 710-106
63.7032 | 0.13736 | 63.6595 | 63.747 | Micro-Vickers | relative_percent | 111 | Micro-Vickers¦SN=709-293¦scale=HV 2¦F=19.62N¦mag=50X | 50X | HV 2 | 2 | 19.62 |  | 709-293
63.2482 | -0.021344 | 63.255 | 63.2415 | Micro-Vickers | relative_percent | 112 | Micro-Vickers¦SN=709-293¦scale=HV 2¦F=19.62N¦mag=50X | 50X | HV 2 | 2 | 19.62 |  | 709-293
62.6977 | -0.51756 | 62.86 | 62.5355 | Micro-Vickers | relative_percent | 113 | Micro-Vickers¦SN=709-293¦scale=HV 2¦F=19.62N¦mag=50X | 50X | HV 2 | 2 | 19.62 |  | 709-293
3.7909 | 0.0036798 | 44.215 | 44.378 | Micro-Vickers | log_ratio | 60 | Micro-Vickers¦SN=710-404¦scale=HV 0.05¦F=0.4905N¦mag=50X | 50X | HV 0.05 | 0.05 | 0.4905 |  | 710-404
3.7761 | 0.0019819 | 43.6025 | 43.689 | Micro-Vickers | log_ratio | 61 | Micro-Vickers¦SN=710-404¦scale=HV 0.05¦F=0.4905N¦mag=50X | 50X | HV 0.05 | 0.05 | 0.4905 |  | 710-404
3.7441 | 0.0037852 | 42.19 | 42.35 | Micro-Vickers | log_ratio | 62 | Micro-Vickers¦SN=710-404¦scale=HV 0.05¦F=0.4905N¦mag=50X | 50X | HV 0.05 | 0.05 | 0.4905 |  | 710-404
3.3537 | 0.0011185 | 28.5925 | 28.6245 | Micro-Vickers | log_ratio | 63 | Micro-Vickers¦SN=711-081¦scale=HV 0.05¦F=0.4905N¦mag=50X | 50X | HV 0.05 | 0.05 | 0.4905 |  | 711-081
3.3717 | 0.0031584 | 29.0825 | 29.1745 | Micro-Vickers | log_ratio | 64 | Micro-Vickers¦SN=711-081¦scale=HV 0.05¦F=0.4905N¦mag=50X | 50X | HV 0.05 | 0.05 | 0.4905 |  | 711-081
3.3553 | 0.0026348 | 28.6175 | 28.693 | Micro-Vickers | log_ratio | 65 | Micro-Vickers¦SN=711-081¦scale=HV 0.05¦F=0.4905N¦mag=50X | 50X | HV 0.05 | 0.05 | 0.4905 |  | 711-081
3.0385 | 0.0029703 | 20.842 | 20.904 | Micro-Vickers | log_ratio | 66 | Micro-Vickers¦SN=710-655¦scale=HV 0.05¦F=0.4905N¦mag=50X | 50X | HV 0.05 | 0.05 | 0.4905 |  | 710-655
3.0324 | -0.0014701 | 20.7625 | 20.732 | Micro-Vickers | log_ratio | 67 | Micro-Vickers¦SN=710-655¦scale=HV 0.05¦F=0.4905N¦mag=50X | 50X | HV 0.05 | 0.05 | 0.4905 |  | 710-655
3.0437 | 0.0037175 | 20.943 | 21.021 | Micro-Vickers | log_ratio | 68 | Micro-Vickers¦SN=710-655¦scale=HV 0.05¦F=0.4905N¦mag=50X | 50X | HV 0.05 | 0.05 | 0.4905 |  | 710-655
4.5481 | 0.0016199 | 94.375 | 94.528 | Micro-Vickers | log_ratio | 69 | Micro-Vickers¦SN=711-081¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 711-081
4.5351 | 0.0025742 | 93.1135 | 93.3535 | Micro-Vickers | log_ratio | 70 | Micro-Vickers¦SN=711-081¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 711-081
4.5457 | 0.00040859 | 94.2065 | 94.245 | Micro-Vickers | log_ratio | 71 | Micro-Vickers¦SN=711-081¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 711-081
3.9882 | -0.0021221 | 54.0125 | 53.898 | Micro-Vickers | log_ratio | 72 | Micro-Vickers¦SN=706-437¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 706-437
3.9851 | 0.0019891 | 53.739 | 53.846 | Micro-Vickers | log_ratio | 73 | Micro-Vickers¦SN=706-437¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 706-437
3.9925 | 0.002602 | 54.119 | 54.26 | Micro-Vickers | log_ratio | 74 | Micro-Vickers¦SN=706-437¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 706-437
3.7605 | 0.00377 | 42.89 | 43.052 | Micro-Vickers | log_ratio | 75 | Micro-Vickers¦SN=711-246¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 711-246
3.7468 | -0.0017576 | 42.424 | 42.3495 | Micro-Vickers | log_ratio | 76 | Micro-Vickers¦SN=711-246¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 711-246
3.7551 | 0.011383 | 42.495 | 42.9815 | Micro-Vickers | log_ratio | 77 | Micro-Vickers¦SN=711-246¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 711-246
3.5575 | -0.0078829 | 35.2145 | 34.938 | Micro-Vickers | log_ratio | 78 | Micro-Vickers¦SN=710-106¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 710-106
3.563 | 0.0025801 | 35.225 | 35.316 | Micro-Vickers | log_ratio | 79 | Micro-Vickers¦SN=710-106¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 710-106
3.5506 | -0.0044068 | 34.9095 | 34.756 | Micro-Vickers | log_ratio | 80 | Micro-Vickers¦SN=710-106¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 710-106
3.4297 | 0.0041792 | 30.8025 | 30.9315 | Micro-Vickers | log_ratio | 81 | Micro-Vickers¦SN=709-293¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 709-293
3.4273 | 0.0017698 | 30.7675 | 30.822 | Micro-Vickers | log_ratio | 82 | Micro-Vickers¦SN=709-293¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 709-293
3.4195 | -0.0037474 | 30.612 | 30.4975 | Micro-Vickers | log_ratio | 83 | Micro-Vickers¦SN=709-293¦scale=HV 0.5¦F=4.905N¦mag=50X | 50X | HV 0.5 | 0.5 | 4.905 |  | 709-293
4.8982 | 0.0020141 | 133.918 | 134.188 | Micro-Vickers | log_ratio | 84 | Micro-Vickers¦SN=711-081¦scale=HV 1¦F=9.81N¦mag=20X | 20X | HV 1 | 1 | 9.81 |  | 711-081
4.8999 | 0.00081179 | 134.217 | 134.326 | Micro-Vickers | log_ratio | 85 | Micro-Vickers¦SN=711-081¦scale=HV 1¦F=9.81N¦mag=20X | 20X | HV 1 | 1 | 9.81 |  | 711-081
4.8938 | 2.6226e-05 | 133.4545 | 133.458 | Micro-Vickers | log_ratio | 86 | Micro-Vickers¦SN=711-081¦scale=HV 1¦F=9.81N¦mag=20X | 20X | HV 1 | 1 | 9.81 |  | 711-081
4.3519 | -0.002744 | 77.7295 | 77.5165 | Micro-Vickers | log_ratio | 87 | Micro-Vickers¦SN=706-437¦scale=HV 1¦F=9.81N¦mag=50X | 50X | HV 1 | 1 | 9.81 |  | 706-437
4.3234 | -0.0059581 | 75.6685 | 75.219 | Micro-Vickers | log_ratio | 88 | Micro-Vickers¦SN=706-437¦scale=HV 1¦F=9.81N¦mag=50X | 50X | HV 1 | 1 | 9.81 |  | 706-437
4.3362 | -0.0055222 | 76.63 | 76.208 | Micro-Vickers | log_ratio | 89 | Micro-Vickers¦SN=706-437¦scale=HV 1¦F=9.81N¦mag=50X | 50X | HV 1 | 1 | 9.81 |  | 706-437
4.1094 | 0.00033655 | 60.902 | 60.9225 | Micro-Vickers | log_ratio | 90 | Micro-Vickers¦SN=711-246¦scale=HV 1¦F=9.81N¦mag=50X | 50X | HV 1 | 1 | 9.81 |  | 711-246
4.0986 | -0.010456 | 60.5695 | 59.9395 | Micro-Vickers | log_ratio | 91 | Micro-Vickers¦SN=711-246¦scale=HV 1¦F=9.81N¦mag=50X | 50X | HV 1 | 1 | 9.81 |  | 711-246
4.0948 | -0.0018909 | 60.0825 | 59.969 | Micro-Vickers | log_ratio | 92 | Micro-Vickers¦SN=711-246¦scale=HV 1¦F=9.81N¦mag=50X | 50X | HV 1 | 1 | 9.81 |  | 711-246
3.9217 | -0.0071801 | 50.668 | 50.3055 | Micro-Vickers | log_ratio | 93 | Micro-Vickers¦SN=710-106¦scale=HV 1¦F=9.81N¦mag=50X | 50X | HV 1 | 1 | 9.81 |  | 710-106
3.916 | -0.0055581 | 50.3365 | 50.0575 | Micro-Vickers | log_ratio | 94 | Micro-Vickers¦SN=710-106¦scale=HV 1¦F=9.81N¦mag=50X | 50X | HV 1 | 1 | 9.81 |  | 710-106
3.9096 | -0.0066361 | 50.0445 | 49.7135 | Micro-Vickers | log_ratio | 95 | Micro-Vickers¦SN=710-106¦scale=HV 1¦F=9.81N¦mag=50X | 50X | HV 1 | 1 | 9.81 |  | 710-106
3.7998 | -0.018123 | 45.101 | 44.291 | Micro-Vickers | log_ratio | 96 | Micro-Vickers¦SN=709-293¦scale=HV 1¦F=9.81N¦mag=50X | 50X | HV 1 | 1 | 9.81 |  | 709-293
3.7975 | -0.018917 | 45.0115 | 44.168 | Micro-Vickers | log_ratio | 97 | Micro-Vickers¦SN=709-293¦scale=HV 1¦F=9.81N¦mag=50X | 50X | HV 1 | 1 | 9.81 |  | 709-293
3.8057 | -0.0116 | 45.22 | 44.6985 | Micro-Vickers | log_ratio | 98 | Micro-Vickers¦SN=709-293¦scale=HV 1¦F=9.81N¦mag=50X | 50X | HV 1 | 1 | 9.81 |  | 709-293
5.2473 | 0.0010629 | 189.9425 | 190.1445 | Micro-Vickers | log_ratio | 99 | Micro-Vickers¦SN=711-081¦scale=HV 2¦F=19.62N¦mag=20X | 20X | HV 2 | 2 | 19.62 |  | 711-081
5.2354 | 0.0022603 | 187.597 | 188.0215 | Micro-Vickers | log_ratio | 100 | Micro-Vickers¦SN=711-081¦scale=HV 2¦F=19.62N¦mag=20X | 20X | HV 2 | 2 | 19.62 |  | 711-081
5.2502 | 0.0021064 | 190.4075 | 190.809 | Micro-Vickers | log_ratio | 101 | Micro-Vickers¦SN=711-081¦scale=HV 2¦F=19.62N¦mag=20X | 20X | HV 2 | 2 | 19.62 |  | 711-081
4.6876 | 0.0004374 | 108.5715 | 108.619 | Micro-Vickers | log_ratio | 102 | Micro-Vickers¦SN=706-437¦scale=HV 2¦F=19.62N¦mag=20X | 20X | HV 2 | 2 | 19.62 |  | 706-437
4.6815 | -0.00055592 | 107.96 | 107.9 | Micro-Vickers | log_ratio | 103 | Micro-Vickers¦SN=706-437¦scale=HV 2¦F=19.62N¦mag=20X | 20X | HV 2 | 2 | 19.62 |  | 706-437
4.6791 | 0.00084977 | 107.63 | 107.7215 | Micro-Vickers | log_ratio | 104 | Micro-Vickers¦SN=706-437¦scale=HV 2¦F=19.62N¦mag=20X | 20X | HV 2 | 2 | 19.62 |  | 706-437
4.4381 | -0.0012881 | 84.672 | 84.563 | Micro-Vickers | log_ratio | 105 | Micro-Vickers¦SN=711-246¦scale=HV 2¦F=19.62N¦mag=50X | 50X | HV 2 | 2 | 19.62 |  | 711-246
4.4429 | -0.00070568 | 85.054 | 84.994 | Micro-Vickers | log_ratio | 106 | Micro-Vickers¦SN=711-246¦scale=HV 2¦F=19.62N¦mag=50X | 50X | HV 2 | 2 | 19.62 |  | 711-246
4.4432 | -0.00021753 | 85.054 | 85.0355 | Micro-Vickers | log_ratio | 107 | Micro-Vickers¦SN=711-246¦scale=HV 2¦F=19.62N¦mag=50X | 50X | HV 2 | 2 | 19.62 |  | 711-246
4.2596 | -0.00058632 | 70.8015 | 70.76 | Micro-Vickers | log_ratio | 108 | Micro-Vickers¦SN=710-106¦scale=HV 2¦F=19.62N¦mag=50X | 50X | HV 2 | 2 | 19.62 |  | 710-106
4.2586 | -0.0049003 | 70.8835 | 70.537 | Micro-Vickers | log_ratio | 109 | Micro-Vickers¦SN=710-106¦scale=HV 2¦F=19.62N¦mag=50X | 50X | HV 2 | 2 | 19.62 |  | 710-106
4.2564 | -0.0050103 | 70.732 | 70.3785 | Micro-Vickers | log_ratio | 110 | Micro-Vickers¦SN=710-106¦scale=HV 2¦F=19.62N¦mag=50X | 50X | HV 2 | 2 | 19.62 |  | 710-106
4.1542 | 0.0013736 | 63.6595 | 63.747 | Micro-Vickers | log_ratio | 111 | Micro-Vickers¦SN=709-293¦scale=HV 2¦F=19.62N¦mag=50X | 50X | HV 2 | 2 | 19.62 |  | 709-293
4.1471 | -0.00021344 | 63.255 | 63.2415 | Micro-Vickers | log_ratio | 112 | Micro-Vickers¦SN=709-293¦scale=HV 2¦F=19.62N¦mag=50X | 50X | HV 2 | 2 | 19.62 |  | 709-293
4.1383 | -0.0051756 | 62.86 | 62.5355 | Micro-Vickers | log_ratio | 113 | Micro-Vickers¦SN=709-293¦scale=HV 2¦F=19.62N¦mag=50X | 50X | HV 2 | 2 | 19.62 |  | 709-293
184.1415 | 0.436 | 183.9235 | 184.3595 | Vickers | absolute | 0 | Vickers¦SN=293-250¦scale=HV 3¦F=29.43N¦mag=20X | 20X | HV 3 | 3 | 29.43 |  | 293-250
184.9338 | -0.4435 | 185.1555 | 184.712 | Vickers | absolute | 1 | Vickers¦SN=293-250¦scale=HV 3¦F=29.43N¦mag=20X | 20X | HV 3 | 3 | 29.43 |  | 293-250
184.6583 | 0.1375 | 184.5895 | 184.727 | Vickers | absolute | 2 | Vickers¦SN=293-250¦scale=HV 3¦F=29.43N¦mag=20X | 20X | HV 3 | 3 | 29.43 |  | 293-250
135.4025 | 0.09 | 135.3575 | 135.4475 | Vickers | absolute | 3 | Vickers¦SN=294-367¦scale=HV 3¦F=29.43N¦mag=20X | 20X | HV 3 | 3 | 29.43 |  | 294-367
134.7893 | -0.0005 | 134.7895 | 134.789 | Vickers | absolute | 4 | Vickers¦SN=294-367¦scale=HV 3¦F=29.43N¦mag=20X | 20X | HV 3 | 3 | 29.43 |  | 294-367
135.0375 | -0.003 | 135.039 | 135.036 | Vickers | absolute | 5 | Vickers¦SN=294-367¦scale=HV 3¦F=29.43N¦mag=20X | 20X | HV 3 | 3 | 29.43 |  | 294-367
103.9393 | 0.3735 | 103.7525 | 104.126 | Vickers | absolute | 6 | Vickers¦SN=294-481¦scale=HV 3¦F=29.43N¦mag=20X | 20X | HV 3 | 3 | 29.43 |  | 294-481
103.9352 | 0.2795 | 103.7955 | 104.075 | Vickers | absolute | 7 | Vickers¦SN=294-481¦scale=HV 3¦F=29.43N¦mag=20X | 20X | HV 3 | 3 | 29.43 |  | 294-481
104.2772 | 0.4125 | 104.071 | 104.4835 | Vickers | absolute | 8 | Vickers¦SN=294-481¦scale=HV 3¦F=29.43N¦mag=20X | 20X | HV 3 | 3 | 29.43 |  | 294-481
86.8935 | -0.526 | 87.1565 | 86.6305 | Vickers | absolute | 9 | Vickers¦SN=216-610¦scale=HV 3¦F=29.43N¦mag=50X | 50X | HV 3 | 3 | 29.43 |  | 216-610
87.983 | 0.356 | 87.805 | 88.161 | Vickers | absolute | 10 | Vickers¦SN=216-610¦scale=HV 3¦F=29.43N¦mag=50X | 50X | HV 3 | 3 | 29.43 |  | 216-610
87.17 | -0.673 | 87.5065 | 86.8335 | Vickers | absolute | 11 | Vickers¦SN=216-610¦scale=HV 3¦F=29.43N¦mag=50X | 50X | HV 3 | 3 | 29.43 |  | 216-610
78.1238 | -0.2365 | 78.242 | 78.0055 | Vickers | absolute | 12 | Vickers¦SN=221-101¦scale=HV 3¦F=29.43N¦mag=50X | 50X | HV 3 | 3 | 29.43 |  | 221-101
77.859 | -0.642 | 78.18 | 77.538 | Vickers | absolute | 13 | Vickers¦SN=221-101¦scale=HV 3¦F=29.43N¦mag=50X | 50X | HV 3 | 3 | 29.43 |  | 221-101
78.2405 | 0.279 | 78.101 | 78.38 | Vickers | absolute | 14 | Vickers¦SN=221-101¦scale=HV 3¦F=29.43N¦mag=50X | 50X | HV 3 | 3 | 29.43 |  | 221-101
478.0763 | -0.7165 | 478.4345 | 477.718 | Vickers | absolute | 15 | Vickers¦SN=293-250¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 293-250
481.7577 | 0.6815 | 481.417 | 482.0985 | Vickers | absolute | 16 | Vickers¦SN=293-250¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 293-250
478.0758 | 0.1795 | 477.986 | 478.1655 | Vickers | absolute | 17 | Vickers¦SN=293-250¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 293-250
349.5155 | 0.554 | 349.2385 | 349.7925 | Vickers | absolute | 18 | Vickers¦SN=294-367¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 294-367
349.5453 | 0.0275 | 349.5315 | 349.559 | Vickers | absolute | 19 | Vickers¦SN=294-367¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 294-367
348.946 | 0.256 | 348.818 | 349.074 | Vickers | absolute | 20 | Vickers¦SN=294-367¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 294-367
269.9153 | -0.0155 | 269.923 | 269.9075 | Vickers | absolute | 21 | Vickers¦SN=294-481¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 294-481
269.748 | -0.227 | 269.8615 | 269.6345 | Vickers | absolute | 22 | Vickers¦SN=294-481¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 294-481
270.3575 | -0.289 | 270.502 | 270.213 | Vickers | absolute | 23 | Vickers¦SN=294-481¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 294-481
227.7202 | -0.0025 | 227.7215 | 227.719 | Vickers | absolute | 24 | Vickers¦SN=216-610¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 216-610
227.3155 | -0.658 | 227.6445 | 226.9865 | Vickers | absolute | 25 | Vickers¦SN=216-610¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 216-610
226.8995 | -0.191 | 226.995 | 226.804 | Vickers | absolute | 26 | Vickers¦SN=216-610¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 216-610
202.724 | 0.074 | 202.687 | 202.761 | Vickers | absolute | 27 | Vickers¦SN=221-101¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 221-101
202.9278 | -0.2555 | 203.0555 | 202.8 | Vickers | absolute | 28 | Vickers¦SN=221-101¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 221-101
202.4298 | -0.5855 | 202.7225 | 202.137 | Vickers | absolute | 29 | Vickers¦SN=221-101¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 221-101
756.7703 | 0.5985 | 756.471 | 757.0695 | Vickers | absolute | 30 | Vickers¦SN=293-250¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 293-250
759.0805 | 0.404 | 758.8785 | 759.2825 | Vickers | absolute | 31 | Vickers¦SN=293-250¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 293-250
753.503 | 0.16 | 753.423 | 753.583 | Vickers | absolute | 32 | Vickers¦SN=293-250¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 293-250
551.7463 | 0.4675 | 551.5125 | 551.98 | Vickers | absolute | 33 | Vickers¦SN=294-367¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 294-367
552.0903 | -0.2085 | 552.1945 | 551.986 | Vickers | absolute | 34 | Vickers¦SN=294-367¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 294-367
551.3025 | -0.141 | 551.373 | 551.232 | Vickers | absolute | 35 | Vickers¦SN=294-367¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 294-367
426.4128 | 0.1755 | 426.325 | 426.5005 | Vickers | absolute | 36 | Vickers¦SN=294-481¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 294-481
426.9595 | 0.499 | 426.71 | 427.209 | Vickers | absolute | 37 | Vickers¦SN=294-481¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 294-481
426.6562 | 0.1595 | 426.5765 | 426.736 | Vickers | absolute | 38 | Vickers¦SN=294-481¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 294-481
360.298 | -0.174 | 360.385 | 360.211 | Vickers | absolute | 39 | Vickers¦SN=216-610¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 216-610
360.6465 | -0.217 | 360.755 | 360.538 | Vickers | absolute | 40 | Vickers¦SN=216-610¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 216-610
360.9525 | -0.439 | 361.172 | 360.733 | Vickers | absolute | 41 | Vickers¦SN=216-610¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 216-610
321.3145 | -0.891 | 321.76 | 320.869 | Vickers | absolute | 42 | Vickers¦SN=221-101¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 221-101
321.6943 | -0.5075 | 321.948 | 321.4405 | Vickers | absolute | 43 | Vickers¦SN=221-101¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 221-101
321.57 | -0.264 | 321.702 | 321.438 | Vickers | absolute | 44 | Vickers¦SN=221-101¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 221-101
1069.108 | 0.329 | 1068.9435 | 1069.2725 | Vickers | absolute | 45 | Vickers¦SN=293-250¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 293-250
1072.5722 | 0.2535 | 1072.4455 | 1072.699 | Vickers | absolute | 46 | Vickers¦SN=293-250¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 293-250
1066.46 | 0.281 | 1066.3195 | 1066.6005 | Vickers | absolute | 47 | Vickers¦SN=293-250¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 293-250
781.8565 | 0.14 | 781.7865 | 781.9265 | Vickers | absolute | 48 | Vickers¦SN=294-367¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 294-367
781.5283 | 0.4215 | 781.3175 | 781.739 | Vickers | absolute | 49 | Vickers¦SN=294-367¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 294-367
781.3925 | 0.105 | 781.34 | 781.445 | Vickers | absolute | 50 | Vickers¦SN=294-367¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 294-367
605.8285 | -0.202 | 605.9295 | 605.7275 | Vickers | absolute | 51 | Vickers¦SN=294-481¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 294-481
606.3255 | 0.079 | 606.286 | 606.365 | Vickers | absolute | 52 | Vickers¦SN=294-481¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 294-481
604.7307 | -0.3775 | 604.9195 | 604.542 | Vickers | absolute | 53 | Vickers¦SN=294-481¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 294-481
511.3195 | -0.055 | 511.347 | 511.292 | Vickers | absolute | 54 | Vickers¦SN=216-610¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 216-610
510.7462 | -0.2075 | 510.85 | 510.6425 | Vickers | absolute | 55 | Vickers¦SN=216-610¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 216-610
511.4167 | -0.1685 | 511.501 | 511.3325 | Vickers | absolute | 56 | Vickers¦SN=216-610¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 216-610
456.064 | 0.041 | 456.0435 | 456.0845 | Vickers | absolute | 57 | Vickers¦SN=221-101¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 221-101
455.8342 | -0.1695 | 455.919 | 455.7495 | Vickers | absolute | 58 | Vickers¦SN=221-101¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 221-101
456.0795 | -0.321 | 456.24 | 455.919 | Vickers | absolute | 59 | Vickers¦SN=221-101¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 221-101
184.1415 | 0.23677 | 183.9235 | 184.3595 | Vickers | relative_percent | 0 | Vickers¦SN=293-250¦scale=HV 3¦F=29.43N¦mag=20X | 20X | HV 3 | 3 | 29.43 |  | 293-250
184.9338 | -0.23982 | 185.1555 | 184.712 | Vickers | relative_percent | 1 | Vickers¦SN=293-250¦scale=HV 3¦F=29.43N¦mag=20X | 20X | HV 3 | 3 | 29.43 |  | 293-250
184.6583 | 0.074462 | 184.5895 | 184.727 | Vickers | relative_percent | 2 | Vickers¦SN=293-250¦scale=HV 3¦F=29.43N¦mag=20X | 20X | HV 3 | 3 | 29.43 |  | 293-250
135.4025 | 0.066468 | 135.3575 | 135.4475 | Vickers | relative_percent | 3 | Vickers¦SN=294-367¦scale=HV 3¦F=29.43N¦mag=20X | 20X | HV 3 | 3 | 29.43 |  | 294-367
134.7893 | -0.00037095 | 134.7895 | 134.789 | Vickers | relative_percent | 4 | Vickers¦SN=294-367¦scale=HV 3¦F=29.43N¦mag=20X | 20X | HV 3 | 3 | 29.43 |  | 294-367
135.0375 | -0.0022216 | 135.039 | 135.036 | Vickers | relative_percent | 5 | Vickers¦SN=294-367¦scale=HV 3¦F=29.43N¦mag=20X | 20X | HV 3 | 3 | 29.43 |  | 294-367
103.9393 | 0.35934 | 103.7525 | 104.126 | Vickers | relative_percent | 6 | Vickers¦SN=294-481¦scale=HV 3¦F=29.43N¦mag=20X | 20X | HV 3 | 3 | 29.43 |  | 294-481
103.9352 | 0.26892 | 103.7955 | 104.075 | Vickers | relative_percent | 7 | Vickers¦SN=294-481¦scale=HV 3¦F=29.43N¦mag=20X | 20X | HV 3 | 3 | 29.43 |  | 294-481
104.2772 | 0.39558 | 104.071 | 104.4835 | Vickers | relative_percent | 8 | Vickers¦SN=294-481¦scale=HV 3¦F=29.43N¦mag=20X | 20X | HV 3 | 3 | 29.43 |  | 294-481
86.8935 | -0.60534 | 87.1565 | 86.6305 | Vickers | relative_percent | 9 | Vickers¦SN=216-610¦scale=HV 3¦F=29.43N¦mag=50X | 50X | HV 3 | 3 | 29.43 |  | 216-610
87.983 | 0.40462 | 87.805 | 88.161 | Vickers | relative_percent | 10 | Vickers¦SN=216-610¦scale=HV 3¦F=29.43N¦mag=50X | 50X | HV 3 | 3 | 29.43 |  | 216-610
87.17 | -0.77205 | 87.5065 | 86.8335 | Vickers | relative_percent | 11 | Vickers¦SN=216-610¦scale=HV 3¦F=29.43N¦mag=50X | 50X | HV 3 | 3 | 29.43 |  | 216-610
78.1238 | -0.30272 | 78.242 | 78.0055 | Vickers | relative_percent | 12 | Vickers¦SN=221-101¦scale=HV 3¦F=29.43N¦mag=50X | 50X | HV 3 | 3 | 29.43 |  | 221-101
77.859 | -0.82457 | 78.18 | 77.538 | Vickers | relative_percent | 13 | Vickers¦SN=221-101¦scale=HV 3¦F=29.43N¦mag=50X | 50X | HV 3 | 3 | 29.43 |  | 221-101
78.2405 | 0.35659 | 78.101 | 78.38 | Vickers | relative_percent | 14 | Vickers¦SN=221-101¦scale=HV 3¦F=29.43N¦mag=50X | 50X | HV 3 | 3 | 29.43 |  | 221-101
478.0763 | -0.14987 | 478.4345 | 477.718 | Vickers | relative_percent | 15 | Vickers¦SN=293-250¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 293-250
481.7577 | 0.14146 | 481.417 | 482.0985 | Vickers | relative_percent | 16 | Vickers¦SN=293-250¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 293-250
478.0758 | 0.037546 | 477.986 | 478.1655 | Vickers | relative_percent | 17 | Vickers¦SN=293-250¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 293-250
349.5155 | 0.15851 | 349.2385 | 349.7925 | Vickers | relative_percent | 18 | Vickers¦SN=294-367¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 294-367
349.5453 | 0.0078674 | 349.5315 | 349.559 | Vickers | relative_percent | 19 | Vickers¦SN=294-367¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 294-367
348.946 | 0.073364 | 348.818 | 349.074 | Vickers | relative_percent | 20 | Vickers¦SN=294-367¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 294-367
269.9153 | -0.0057425 | 269.923 | 269.9075 | Vickers | relative_percent | 21 | Vickers¦SN=294-481¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 294-481
269.748 | -0.084153 | 269.8615 | 269.6345 | Vickers | relative_percent | 22 | Vickers¦SN=294-481¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 294-481
270.3575 | -0.1069 | 270.502 | 270.213 | Vickers | relative_percent | 23 | Vickers¦SN=294-481¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 294-481
227.7202 | -0.0010978 | 227.7215 | 227.719 | Vickers | relative_percent | 24 | Vickers¦SN=216-610¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 216-610
227.3155 | -0.28947 | 227.6445 | 226.9865 | Vickers | relative_percent | 25 | Vickers¦SN=216-610¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 216-610
226.8995 | -0.084178 | 226.995 | 226.804 | Vickers | relative_percent | 26 | Vickers¦SN=216-610¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 216-610
202.724 | 0.036503 | 202.687 | 202.761 | Vickers | relative_percent | 27 | Vickers¦SN=221-101¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 221-101
202.9278 | -0.12591 | 203.0555 | 202.8 | Vickers | relative_percent | 28 | Vickers¦SN=221-101¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 221-101
202.4298 | -0.28924 | 202.7225 | 202.137 | Vickers | relative_percent | 29 | Vickers¦SN=221-101¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 221-101
756.7703 | 0.079086 | 756.471 | 757.0695 | Vickers | relative_percent | 30 | Vickers¦SN=293-250¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 293-250
759.0805 | 0.053222 | 758.8785 | 759.2825 | Vickers | relative_percent | 31 | Vickers¦SN=293-250¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 293-250
753.503 | 0.021234 | 753.423 | 753.583 | Vickers | relative_percent | 32 | Vickers¦SN=293-250¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 293-250
551.7463 | 0.084731 | 551.5125 | 551.98 | Vickers | relative_percent | 33 | Vickers¦SN=294-367¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 294-367
552.0903 | -0.037766 | 552.1945 | 551.986 | Vickers | relative_percent | 34 | Vickers¦SN=294-367¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 294-367
551.3025 | -0.025576 | 551.373 | 551.232 | Vickers | relative_percent | 35 | Vickers¦SN=294-367¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 294-367
426.4128 | 0.041157 | 426.325 | 426.5005 | Vickers | relative_percent | 36 | Vickers¦SN=294-481¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 294-481
426.9595 | 0.11687 | 426.71 | 427.209 | Vickers | relative_percent | 37 | Vickers¦SN=294-481¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 294-481
426.6562 | 0.037384 | 426.5765 | 426.736 | Vickers | relative_percent | 38 | Vickers¦SN=294-481¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 294-481
360.298 | -0.048293 | 360.385 | 360.211 | Vickers | relative_percent | 39 | Vickers¦SN=216-610¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 216-610
360.6465 | -0.06017 | 360.755 | 360.538 | Vickers | relative_percent | 40 | Vickers¦SN=216-610¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 216-610
360.9525 | -0.12162 | 361.172 | 360.733 | Vickers | relative_percent | 41 | Vickers¦SN=216-610¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 216-610
321.3145 | -0.2773 | 321.76 | 320.869 | Vickers | relative_percent | 42 | Vickers¦SN=221-101¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 221-101
321.6943 | -0.15776 | 321.948 | 321.4405 | Vickers | relative_percent | 43 | Vickers¦SN=221-101¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 221-101
321.57 | -0.082097 | 321.702 | 321.438 | Vickers | relative_percent | 44 | Vickers¦SN=221-101¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 221-101
1069.108 | 0.030773 | 1068.9435 | 1069.2725 | Vickers | relative_percent | 45 | Vickers¦SN=293-250¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 293-250
1072.5722 | 0.023635 | 1072.4455 | 1072.699 | Vickers | relative_percent | 46 | Vickers¦SN=293-250¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 293-250
1066.46 | 0.026349 | 1066.3195 | 1066.6005 | Vickers | relative_percent | 47 | Vickers¦SN=293-250¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 293-250
781.8565 | 0.017906 | 781.7865 | 781.9265 | Vickers | relative_percent | 48 | Vickers¦SN=294-367¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 294-367
781.5283 | 0.053933 | 781.3175 | 781.739 | Vickers | relative_percent | 49 | Vickers¦SN=294-367¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 294-367
781.3925 | 0.013438 | 781.34 | 781.445 | Vickers | relative_percent | 50 | Vickers¦SN=294-367¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 294-367
605.8285 | -0.033343 | 605.9295 | 605.7275 | Vickers | relative_percent | 51 | Vickers¦SN=294-481¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 294-481
606.3255 | 0.013029 | 606.286 | 606.365 | Vickers | relative_percent | 52 | Vickers¦SN=294-481¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 294-481
604.7307 | -0.062424 | 604.9195 | 604.542 | Vickers | relative_percent | 53 | Vickers¦SN=294-481¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 294-481
511.3195 | -0.010756 | 511.347 | 511.292 | Vickers | relative_percent | 54 | Vickers¦SN=216-610¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 216-610
510.7462 | -0.040627 | 510.85 | 510.6425 | Vickers | relative_percent | 55 | Vickers¦SN=216-610¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 216-610
511.4167 | -0.032948 | 511.501 | 511.3325 | Vickers | relative_percent | 56 | Vickers¦SN=216-610¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 216-610
456.064 | 0.00899 | 456.0435 | 456.0845 | Vickers | relative_percent | 57 | Vickers¦SN=221-101¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 221-101
455.8342 | -0.037185 | 455.919 | 455.7495 | Vickers | relative_percent | 58 | Vickers¦SN=221-101¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 221-101
456.0795 | -0.070382 | 456.24 | 455.919 | Vickers | relative_percent | 59 | Vickers¦SN=221-101¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 221-101
5.2157 | 0.0023677 | 183.9235 | 184.3595 | Vickers | log_ratio | 0 | Vickers¦SN=293-250¦scale=HV 3¦F=29.43N¦mag=20X | 20X | HV 3 | 3 | 29.43 |  | 293-250
5.22 | -0.0023982 | 185.1555 | 184.712 | Vickers | log_ratio | 1 | Vickers¦SN=293-250¦scale=HV 3¦F=29.43N¦mag=20X | 20X | HV 3 | 3 | 29.43 |  | 293-250
5.2185 | 0.00074462 | 184.5895 | 184.727 | Vickers | log_ratio | 2 | Vickers¦SN=293-250¦scale=HV 3¦F=29.43N¦mag=20X | 20X | HV 3 | 3 | 29.43 |  | 293-250
4.9083 | 0.00066468 | 135.3575 | 135.4475 | Vickers | log_ratio | 3 | Vickers¦SN=294-367¦scale=HV 3¦F=29.43N¦mag=20X | 20X | HV 3 | 3 | 29.43 |  | 294-367
4.9037 | -3.7095e-06 | 134.7895 | 134.789 | Vickers | log_ratio | 4 | Vickers¦SN=294-367¦scale=HV 3¦F=29.43N¦mag=20X | 20X | HV 3 | 3 | 29.43 |  | 294-367
4.9056 | -2.2216e-05 | 135.039 | 135.036 | Vickers | log_ratio | 5 | Vickers¦SN=294-367¦scale=HV 3¦F=29.43N¦mag=20X | 20X | HV 3 | 3 | 29.43 |  | 294-367
4.6438 | 0.0035934 | 103.7525 | 104.126 | Vickers | log_ratio | 6 | Vickers¦SN=294-481¦scale=HV 3¦F=29.43N¦mag=20X | 20X | HV 3 | 3 | 29.43 |  | 294-481
4.6438 | 0.0026892 | 103.7955 | 104.075 | Vickers | log_ratio | 7 | Vickers¦SN=294-481¦scale=HV 3¦F=29.43N¦mag=20X | 20X | HV 3 | 3 | 29.43 |  | 294-481
4.6471 | 0.0039558 | 104.071 | 104.4835 | Vickers | log_ratio | 8 | Vickers¦SN=294-481¦scale=HV 3¦F=29.43N¦mag=20X | 20X | HV 3 | 3 | 29.43 |  | 294-481
4.4647 | -0.0060534 | 87.1565 | 86.6305 | Vickers | log_ratio | 9 | Vickers¦SN=216-610¦scale=HV 3¦F=29.43N¦mag=50X | 50X | HV 3 | 3 | 29.43 |  | 216-610
4.4771 | 0.0040462 | 87.805 | 88.161 | Vickers | log_ratio | 10 | Vickers¦SN=216-610¦scale=HV 3¦F=29.43N¦mag=50X | 50X | HV 3 | 3 | 29.43 |  | 216-610
4.4679 | -0.0077206 | 87.5065 | 86.8335 | Vickers | log_ratio | 11 | Vickers¦SN=216-610¦scale=HV 3¦F=29.43N¦mag=50X | 50X | HV 3 | 3 | 29.43 |  | 216-610
4.3583 | -0.0030273 | 78.242 | 78.0055 | Vickers | log_ratio | 12 | Vickers¦SN=221-101¦scale=HV 3¦F=29.43N¦mag=50X | 50X | HV 3 | 3 | 29.43 |  | 221-101
4.3549 | -0.0082457 | 78.18 | 77.538 | Vickers | log_ratio | 13 | Vickers¦SN=221-101¦scale=HV 3¦F=29.43N¦mag=50X | 50X | HV 3 | 3 | 29.43 |  | 221-101
4.3598 | 0.0035659 | 78.101 | 78.38 | Vickers | log_ratio | 14 | Vickers¦SN=221-101¦scale=HV 3¦F=29.43N¦mag=50X | 50X | HV 3 | 3 | 29.43 |  | 221-101
6.1698 | -0.0014987 | 478.4345 | 477.718 | Vickers | log_ratio | 15 | Vickers¦SN=293-250¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 293-250
6.1774 | 0.0014146 | 481.417 | 482.0985 | Vickers | log_ratio | 16 | Vickers¦SN=293-250¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 293-250
6.1698 | 0.00037546 | 477.986 | 478.1655 | Vickers | log_ratio | 17 | Vickers¦SN=293-250¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 293-250
5.8565 | 0.0015851 | 349.2385 | 349.7925 | Vickers | log_ratio | 18 | Vickers¦SN=294-367¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 294-367
5.8566 | 7.8674e-05 | 349.5315 | 349.559 | Vickers | log_ratio | 19 | Vickers¦SN=294-367¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 294-367
5.8549 | 0.00073364 | 348.818 | 349.074 | Vickers | log_ratio | 20 | Vickers¦SN=294-367¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 294-367
5.5981 | -5.7425e-05 | 269.923 | 269.9075 | Vickers | log_ratio | 21 | Vickers¦SN=294-481¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 294-481
5.5975 | -0.00084153 | 269.8615 | 269.6345 | Vickers | log_ratio | 22 | Vickers¦SN=294-481¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 294-481
5.5997 | -0.001069 | 270.502 | 270.213 | Vickers | log_ratio | 23 | Vickers¦SN=294-481¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 294-481
5.4281 | -1.0978e-05 | 227.7215 | 227.719 | Vickers | log_ratio | 24 | Vickers¦SN=216-610¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 216-610
5.4263 | -0.0028947 | 227.6445 | 226.9865 | Vickers | log_ratio | 25 | Vickers¦SN=216-610¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 216-610
5.4245 | -0.00084178 | 226.995 | 226.804 | Vickers | log_ratio | 26 | Vickers¦SN=216-610¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 216-610
5.3118 | 0.00036503 | 202.687 | 202.761 | Vickers | log_ratio | 27 | Vickers¦SN=221-101¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 221-101
5.3128 | -0.0012591 | 203.0555 | 202.8 | Vickers | log_ratio | 28 | Vickers¦SN=221-101¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 221-101
5.3104 | -0.0028924 | 202.7225 | 202.137 | Vickers | log_ratio | 29 | Vickers¦SN=221-101¦scale=HV 20¦F=196.2N¦mag=20X | 20X | HV 20 | 20 | 196.2 |  | 221-101
6.6291 | 0.00079086 | 756.471 | 757.0695 | Vickers | log_ratio | 30 | Vickers¦SN=293-250¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 293-250
6.6321 | 0.00053222 | 758.8785 | 759.2825 | Vickers | log_ratio | 31 | Vickers¦SN=293-250¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 293-250
6.6247 | 0.00021234 | 753.423 | 753.583 | Vickers | log_ratio | 32 | Vickers¦SN=293-250¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 293-250
6.3131 | 0.00084731 | 551.5125 | 551.98 | Vickers | log_ratio | 33 | Vickers¦SN=294-367¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 294-367
6.3137 | -0.00037766 | 552.1945 | 551.986 | Vickers | log_ratio | 34 | Vickers¦SN=294-367¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 294-367
6.3123 | -0.00025576 | 551.373 | 551.232 | Vickers | log_ratio | 35 | Vickers¦SN=294-367¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 294-367
6.0554 | 0.00041157 | 426.325 | 426.5005 | Vickers | log_ratio | 36 | Vickers¦SN=294-481¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 294-481
6.0567 | 0.0011687 | 426.71 | 427.209 | Vickers | log_ratio | 37 | Vickers¦SN=294-481¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 294-481
6.056 | 0.00037384 | 426.5765 | 426.736 | Vickers | log_ratio | 38 | Vickers¦SN=294-481¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 294-481
5.8869 | -0.00048293 | 360.385 | 360.211 | Vickers | log_ratio | 39 | Vickers¦SN=216-610¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 216-610
5.8879 | -0.0006017 | 360.755 | 360.538 | Vickers | log_ratio | 40 | Vickers¦SN=216-610¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 216-610
5.8887 | -0.0012162 | 361.172 | 360.733 | Vickers | log_ratio | 41 | Vickers¦SN=216-610¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 216-610
5.7724 | -0.002773 | 321.76 | 320.869 | Vickers | log_ratio | 42 | Vickers¦SN=221-101¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 221-101
5.7736 | -0.0015776 | 321.948 | 321.4405 | Vickers | log_ratio | 43 | Vickers¦SN=221-101¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 221-101
5.7732 | -0.00082097 | 321.702 | 321.438 | Vickers | log_ratio | 44 | Vickers¦SN=221-101¦scale=HV 50¦F=490.5N¦mag=20X | 20X | HV 50 | 50 | 490.5 |  | 221-101
6.9746 | 0.00030773 | 1068.9435 | 1069.2725 | Vickers | log_ratio | 45 | Vickers¦SN=293-250¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 293-250
6.9778 | 0.00023635 | 1072.4455 | 1072.699 | Vickers | log_ratio | 46 | Vickers¦SN=293-250¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 293-250
6.9721 | 0.00026349 | 1066.3195 | 1066.6005 | Vickers | log_ratio | 47 | Vickers¦SN=293-250¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 293-250
6.6617 | 0.00017906 | 781.7865 | 781.9265 | Vickers | log_ratio | 48 | Vickers¦SN=294-367¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 294-367
6.6613 | 0.00053933 | 781.3175 | 781.739 | Vickers | log_ratio | 49 | Vickers¦SN=294-367¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 294-367
6.6611 | 0.00013438 | 781.34 | 781.445 | Vickers | log_ratio | 50 | Vickers¦SN=294-367¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 294-367
6.4066 | -0.00033343 | 605.9295 | 605.7275 | Vickers | log_ratio | 51 | Vickers¦SN=294-481¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 294-481
6.4074 | 0.00013029 | 606.286 | 606.365 | Vickers | log_ratio | 52 | Vickers¦SN=294-481¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 294-481
6.4048 | -0.00062424 | 604.9195 | 604.542 | Vickers | log_ratio | 53 | Vickers¦SN=294-481¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 294-481
6.237 | -0.00010756 | 511.347 | 511.292 | Vickers | log_ratio | 54 | Vickers¦SN=216-610¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 216-610
6.2359 | -0.00040627 | 510.85 | 510.6425 | Vickers | log_ratio | 55 | Vickers¦SN=216-610¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 216-610
6.2372 | -0.00032948 | 511.501 | 511.3325 | Vickers | log_ratio | 56 | Vickers¦SN=216-610¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 216-610
6.1226 | 8.99e-05 | 456.0435 | 456.0845 | Vickers | log_ratio | 57 | Vickers¦SN=221-101¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 221-101
6.1221 | -0.00037185 | 455.919 | 455.7495 | Vickers | log_ratio | 58 | Vickers¦SN=221-101¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 221-101
6.1227 | -0.00070382 | 456.24 | 455.919 | Vickers | log_ratio | 59 | Vickers¦SN=221-101¦scale=HV 100¦F=981N¦mag=20X | 20X | HV 100 | 100 | 981 |  | 221-101

## Deming
method | model | lambda_ratio | n | intercept | slope | residual_sd | intercept_ci_low | intercept_ci_high | slope_ci_low | slope_ci_high | measurand
Brinell | unweighted_deming | 0.5 | 36 | 4.8817 | 1.0047 | 13.491 | -1.9772 | 13.0094 | 1.0026 | 1.0161 | individual_dimension
Brinell | unweighted_deming | 1 | 36 | 4.9007 | 1.0047 | 13.4909 | -1.9615 | 13.0551 | 1.0026 | 1.0161 | individual_dimension
Brinell | unweighted_deming | 2 | 36 | 4.9197 | 1.0047 | 13.4908 | -1.9456 | 13.1009 | 1.0026 | 1.0161 | individual_dimension
Micro-Vickers | unweighted_deming | 0.5 | 54 | -0.20445 | 1.0022 | 0.26374 | -0.44529 | 0.054669 | 0.99806 | 1.0042 | individual_dimension
Micro-Vickers | unweighted_deming | 1 | 54 | -0.20399 | 1.0022 | 0.26374 | -0.44446 | 0.055316 | 0.99805 | 1.0042 | individual_dimension
Micro-Vickers | unweighted_deming | 2 | 54 | -0.20352 | 1.0021 | 0.26374 | -0.44377 | 0.055963 | 0.99804 | 1.0042 | individual_dimension
Vickers | unweighted_deming | 0.5 | 60 | -0.19843 | 1.0004 | 0.34967 | -0.38287 | 0.0019191 | 0.99998 | 1.0008 | individual_dimension
Vickers | unweighted_deming | 1 | 60 | -0.19831 | 1.0004 | 0.34967 | -0.38274 | 0.0020815 | 0.99998 | 1.0008 | individual_dimension
Vickers | unweighted_deming | 2 | 60 | -0.19819 | 1.0004 | 0.34967 | -0.38261 | 0.002244 | 0.99998 | 1.0008 | individual_dimension
Brinell | unweighted_deming | 0.5 | 12 | -1.1487 | 0.98817 | 6.2218 | -4.6491 | 7.8857 | 0.94376 | 1.0023 | series_mean_hardness
Brinell | unweighted_deming | 1 | 12 | -1.0706 | 0.98791 | 6.2209 | -4.5689 | 7.9206 | 0.94331 | 1.0022 | series_mean_hardness
Brinell | unweighted_deming | 2 | 12 | -0.99307 | 0.98766 | 6.2202 | -4.5532 | 7.954 | 0.94288 | 1.0022 | series_mean_hardness
Micro-Vickers | unweighted_deming | 0.5 | 18 | -2.77 | 1.0121 | 6.5418 | -7.8533 | -0.10209 | 1.0017 | 1.0271 | series_mean_hardness
Micro-Vickers | unweighted_deming | 1 | 18 | -2.7385 | 1.0121 | 6.5415 | -7.8043 | -0.094748 | 1.0017 | 1.027 | series_mean_hardness
Micro-Vickers | unweighted_deming | 2 | 18 | -2.7068 | 1.012 | 6.5414 | -7.7366 | -0.087409 | 1.0017 | 1.0268 | series_mean_hardness
Vickers | unweighted_deming | 0.5 | 20 | -1.5285 | 1.0042 | 1.4945 | -2.7493 | -0.74436 | 1.002 | 1.0065 | series_mean_hardness
Vickers | unweighted_deming | 1 | 20 | -1.5259 | 1.0042 | 1.4945 | -2.7425 | -0.74338 | 1.002 | 1.0065 | series_mean_hardness
Vickers | unweighted_deming | 2 | 20 | -1.5234 | 1.0042 | 1.4945 | -2.7357 | -0.7424 | 1.002 | 1.0065 | series_mean_hardness

## ODR_Series_Hardness
method | model | status | n | intercept | slope | intercept_sd | slope_sd | residual_variance
Brinell | weighted_odr | run | 12 | -1.441 | 0.99804 | 0.76764 | 0.0036662 | 0.047337
Micro-Vickers | weighted_odr | run | 18 | -1.0026 | 1.007 | 0.51798 | 0.003023 | 0.6863
Vickers | weighted_odr | run | 20 | -0.91989 | 1.0024 | 0.39985 | 0.0010455 | 0.057048

## En_Results
series_id | method | serial_number | hardness_scale_raw | hardness_scale | hardness_level | applied_force_n | ball_diameter_mm | magnification | n_positions | excel_rows | picture_ids | manual_dimension_mean_from_individual_um | manual_dimension_sd_from_individual_um | manual_hardness_mean_from_individual | manual_hardness_sd_from_individual | manual_series_mean_dimension_um | manual_series_mean_hardness | manual_expanded_uncertainty | manual_relative_expanded_uncertainty | automatic_dimension_mean_from_individual_um | automatic_dimension_sd_from_individual_um | automatic_hardness_mean_from_individual | automatic_hardness_sd_from_individual | automatic_series_mean_dimension_um | automatic_series_mean_hardness | automatic_expanded_uncertainty | automatic_relative_expanded_uncertainty | manual_standard_uncertainty | automatic_standard_uncertainty | en_numerator | en_denominator | en | abs_en | en_classification
Brinell¦SN=290-736¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | Brinell | 290-736 | HBW 1/30 | HBW 1/30 | 1 | 294.3 | 1 | 20X | 3 | 18,19,20 |  | 254.9415 | 0.35075 | 578.167 | 1.616 | 254.9415 | 578.167 | 11.6423 | 2.0137 | 255.4053 | 0.97174 | 576.049 | 4.4452 | 255.4053 | 576.049 | 14.0586 | 2.4405 | 5.8211 | 7.0293 | -2.118 | 18.2534 | -0.11603 | 0.11603 | satisfactory screening result
Brinell¦SN=290-736¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | Brinell | 290-736 | HBW 1/5 | HBW 1/5 | 1 | 49.05 | 1 | 20X | 3 | 9,10,11 | 600_HBW_1_5_1_20X | 116.9922 | 1.9329 | 463.921 | 15.1635 | 116.9922 | 463.921 | 19.6775 | 4.2416 | 117.9267 | 1.8837 | 456.5573 | 14.5052 | 117.9267 | 456.5573 | 19.075 | 4.178 | 9.8387 | 9.5375 | -7.3637 | 27.4055 | -0.26869 | 0.26869 | satisfactory screening result
Brinell¦SN=293-250¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | Brinell | 293-250 | HBW 1/30 | HBW 1/30 | 1 | 294.3 | 1 | 20X | 3 | 12,13,14 |  | 464.74 | 0.40473 | 166.7753 | 0.30939 | 464.74 | 166.7753 | 2.7303 | 1.6371 | 466.62 | 0.9328 | 165.3487 | 0.70327 | 466.62 | 165.3487 | 3.4346 | 2.0772 | 1.3651 | 1.7173 | -1.4267 | 4.3876 | -0.32516 | 0.32516 | satisfactory screening result
Brinell¦SN=293-250¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | Brinell | 293-250 | HBW 1/5 | HBW 1/5 | 1 | 49.05 | 1 | 20X | 3 | 3,4,5 | 100_HBW_1_5_1_20X | 210.4625 | 1.173 | 142.1683 | 1.6097 | 210.4625 | 142.1683 | 3.9649 | 2.7889 | 211.8828 | 0.70577 | 140.241 | 0.94754 | 211.8828 | 140.241 | 3.3499 | 2.3886 | 1.9824 | 1.6749 | -1.9273 | 5.1906 | -0.37132 | 0.37132 | satisfactory screening result
Brinell¦SN=294-367¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | Brinell | 294-367 | HBW 1/30 | HBW 1/30 | 1 | 294.3 | 1 | 20X | 3 | 15,16,17 |  | 350.6645 | 1.1464 | 300.8717 | 2.0306 | 350.6645 | 300.8717 | 7.3022 | 2.427 | 351.1102 | 1.2353 | 300.084 | 2.1781 | 351.1102 | 300.084 | 7.5794 | 2.5258 | 3.6511 | 3.7897 | -0.78767 | 10.5247 | -0.07484 | 0.07484 | satisfactory screening result
Brinell¦SN=294-367¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | Brinell | 294-367 | HBW 1/5 | HBW 1/5 | 1 | 49.05 | 1 | 20X | 3 | 6,7,8 | 300_HBW_1_5_1_20X | 154.5825 | 0.18559 | 264.8983 | 0.64014 | 154.5825 | 264.8983 | 6.8031 | 2.5682 | 155.1857 | 0.38871 | 262.832 | 1.3271 | 155.1857 | 262.832 | 6.8719 | 2.6145 | 3.4016 | 3.4359 | -2.0663 | 9.6698 | -0.21369 | 0.21369 | satisfactory screening result
Brinell¦SN=810-202¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | Brinell | 810-202 | HBW 10/3000 | HBW 10/3000 | 10 | 29430 | 10 | 20X | 3 | 33,34,35 |  | 3804.9168 | 3.0006 | 253.998 | 0.41743 | 3804.9168 | 253.998 | 11.9262 | 4.6954 | 3823.7332 | 15.6461 | 251.411 | 2.1386 | 3823.7332 | 251.411 | 60.0155 | 23.8715 | 5.9631 | 30.0078 | -2.587 | 61.1891 | -0.042279 | 0.042279 | satisfactory screening result
Brinell¦SN=810-202¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=20X | Brinell | 810-202 | HBW 10/500 | HBW 10/500 | 10 | 4905 | 10 | 20X | 3 | 24,25,26 |  | 1686.2058 | 1.1722 | 222.3693 | 0.31133 | 1686.2058 | 222.3693 | 4.7157 | 2.1207 | 1710.5187 | 10.5711 | 216.0637 | 2.698 | 1710.5187 | 216.0637 | 34.9003 | 16.1528 | 2.3578 | 17.4502 | -6.3057 | 35.2175 | -0.17905 | 0.17905 | satisfactory screening result
Brinell¦SN=812-006¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | Brinell | 812-006 | HBW 10/3000 | HBW 10/3000 | 10 | 29430 | 10 | 20X | 3 | 30,31,32 |  | 4840.7873 | 9.3405 | 152.8677 | 0.63137 | 4840.7873 | 152.8677 | 21.8225 | 14.2754 | 4861.091 | 16.8025 | 151.5053 | 1.12 | 4861.091 | 151.5053 | 38.8338 | 25.632 | 10.9112 | 19.4169 | -1.3623 | 44.5453 | -0.030583 | 0.030583 | satisfactory screening result
Brinell¦SN=812-006¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=20X | Brinell | 812-006 | HBW 10/500 | HBW 10/500 | 10 | 4905 | 10 | 20X | 3 | 21,22,23 |  | 2203.1337 | 4.495 | 129.5897 | 0.53463 | 2203.1337 | 129.5897 | 8.9941 | 6.9405 | 2234.308 | 5.5375 | 125.954 | 0.6338 | 2234.308 | 125.954 | 10.7219 | 8.5126 | 4.4971 | 5.361 | -3.6357 | 13.9948 | -0.25979 | 0.25979 | satisfactory screening result
Brinell¦SN=8301-50DW¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | Brinell | 8301-50DW | HBW 10/3000 | HBW 10/3000 | 10 | 29430 | 10 | 20X | 3 | 36,37,38 |  | 2649.1663 | 6.4182 | 534.7197 | 2.6333 | 2649.1663 | 534.7197 | 52.634 | 9.8433 | 2655.9918 | 7.5572 | 531.926 | 3.0861 | 2655.9918 | 531.926 | 61.5446 | 11.5701 | 26.317 | 30.7723 | -2.7937 | 80.982 | -0.034497 | 0.034497 | satisfactory screening result
Brinell¦SN=8301-50DW¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=10X | Brinell | 8301-50DW | HBW 10/500 | HBW 10/500 | 10 | 4905 | 10 | 10X | 3 | 27,28,29 |  | 1204.3738 | 1.3301 | 437.4327 | 0.97077 | 1204.3738 | 437.4327 | 10.3037 | 2.3555 | 1239.9103 | 21.3949 | 412.87 | 14.1123 | 1239.9103 | 412.87 | 134.7277 | 32.632 | 5.1519 | 67.3638 | -24.5627 | 135.1211 | -0.18178 | 0.18178 | satisfactory screening result
Micro-Vickers¦SN=706-437¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers | 706-437 | HV 0,5 | HV 0.5 | 0.5 | 4.905 |  | 50X | 3 | 15,16,17 | 706-437_307_HMV_0.5_1_50X | 53.9568 | 0.19602 | 318.5757 | 2.2664 | 53.9568 | 318.5757 | 4.4691 | 1.4029 | 54.0013 | 0.22552 | 317.8967 | 2.6461 | 54.0013 | 317.8967 | 4.4913 | 1.4128 | 2.2346 | 2.2456 | -0.679 | 6.336 | -0.10717 | 0.10717 | satisfactory screening result
Micro-Vickers¦SN=706-437¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers | 706-437 | HV 1 | HV 1 | 1 | 9.81 |  | 50X | 3 | 30,31,32 | 706-437_307_HMV_1_1_50X; 706-437_307_HMV_1_2_50X | 76.676 | 1.0313 | 315.334 | 8.2865 | 76.676 | 315.334 | 6.3542 | 2.0151 | 76.3145 | 1.1524 | 318.487 | 9.5923 | 76.3145 | 318.487 | 6.8887 | 2.1629 | 3.1771 | 3.4444 | 3.153 | 9.3718 | 0.33644 | 0.33644 | satisfactory screening result
Micro-Vickers¦SN=706-437¦scale=HV 2¦F=19.62N¦mag=20X | Micro-Vickers | 706-437 | HV 2 | HV 2 | 2 | 19.62 |  | 20X | 3 | 45,46,47 | 706-437_307_HMV_2_1_50X | 108.0538 | 0.47771 | 317.597 | 2.8033 | 108.0538 | 317.597 | 4.4175 | 1.3909 | 108.0802 | 0.4751 | 317.443 | 2.7831 | 108.0802 | 317.443 | 4.4086 | 1.3888 | 2.2087 | 2.2043 | -0.154 | 6.241 | -0.024676 | 0.024676 | satisfactory screening result
Micro-Vickers¦SN=709-293¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers | 709-293 | HV 0,5 | HV 0.5 | 0.5 | 4.905 |  | 50X | 3 | 24,25,26 | 709-293_904_HMV_0.5_1_50X | 30.7273 | 0.1014 | 981.8373 | 6.4964 | 30.7273 | 981.8373 | 16.2799 | 1.6581 | 30.7503 | 0.2257 | 980.475 | 14.4661 | 30.7503 | 980.475 | 16.5296 | 1.6859 | 8.1399 | 8.2648 | -1.3623 | 23.2005 | -0.05872 | 0.05872 | satisfactory screening result
Micro-Vickers¦SN=709-293¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers | 709-293 | HV 1 | HV 1 | 1 | 9.81 |  | 50X | 3 | 39,40,41 | 709-293_904_HMV_1_1_50X | 45.1108 | 0.1046 | 911.0723 | 4.2229 | 45.1108 | 911.0723 | 13.23 | 1.4521 | 44.3858 | 0.27767 | 941.148 | 11.7137 | 44.3858 | 941.148 | 14.2215 | 1.5111 | 6.615 | 7.1107 | 30.0757 | 19.4237 | 1.5484 | 1.5484 | result requiring investigation
Micro-Vickers¦SN=709-293¦scale=HV 2¦F=19.62N¦mag=50X | Micro-Vickers | 709-293 | HV 2 | HV 2 | 2 | 19.62 |  | 50X | 3 | 54,55,56 | 709-293_904_HMV_2_1_50X | 63.2582 | 0.39976 | 926.7043 | 11.711 | 63.2582 | 926.7043 | 13.4406 | 1.4504 | 63.1747 | 0.60851 | 929.244 | 17.937 | 63.1747 | 929.244 | 14.9652 | 1.6105 | 6.7203 | 7.4826 | 2.5397 | 20.1149 | 0.12626 | 0.12626 | satisfactory screening result
Micro-Vickers¦SN=710-106¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers | 710-106 | HV 0,5 | HV 0.5 | 0.5 | 4.905 |  | 50X | 3 | 21,22,23 | 710-106_708_HMV_0.5_1_50X | 35.1163 | 0.1792 | 751.7683 | 7.7062 | 35.1163 | 751.7683 | 11.9766 | 1.5931 | 35.0033 | 0.28566 | 756.691 | 12.3017 | 35.0033 | 756.691 | 12.3385 | 1.6306 | 5.9883 | 6.1692 | 4.9227 | 17.1952 | 0.28628 | 0.28628 | satisfactory screening result
Micro-Vickers¦SN=710-106¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers | 710-106 | HV 1 | HV 1 | 1 | 9.81 |  | 50X | 3 | 36,37,38 | 710-106_708_HMV_1_1_50X | 50.3497 | 0.31196 | 731.3917 | 9.0579 | 50.3497 | 731.3917 | 10.7947 | 1.4759 | 50.0255 | 0.29729 | 740.9113 | 8.8186 | 50.0255 | 740.9113 | 10.9011 | 1.4713 | 5.3974 | 5.4506 | 9.5197 | 15.3415 | 0.62052 | 0.62052 | satisfactory screening result
Micro-Vickers¦SN=710-106¦scale=HV 2¦F=19.62N¦mag=50X | Micro-Vickers | 710-106 | HV 2 | HV 2 | 2 | 19.62 |  | 50X | 3 | 51,52,53 | 710-106_708_HMV_2_1_50X | 70.8057 | 0.075836 | 739.6133 | 1.5842 | 70.8057 | 739.6133 | 9.5237 | 1.2877 | 70.5585 | 0.19166 | 744.817 | 4.0484 | 70.5585 | 744.817 | 9.8039 | 1.3163 | 4.7619 | 4.9019 | 5.2037 | 13.6681 | 0.38072 | 0.38072 | satisfactory screening result
Micro-Vickers¦SN=710-404¦scale=HV 0.05¦F=0.4905N¦mag=50X | Micro-Vickers | 710-404 | HV 0,5 | HV 0.05 | 0.05 | 0.4905 |  | 50X | 3 | 3,4,5 | 710-404_39.6_HMV_0.05_1_50X | 43.3358 | 1.0385 | 49.4187 | 2.3995 | 43.3358 | 49.4187 | 1.0647 | 2.1544 | 43.4723 | 1.0312 | 49.1073 | 2.3551 | 43.4723 | 49.1073 | 1.0535 | 2.1453 | 0.53235 | 0.52675 | -0.31133 | 1.4978 | -0.20786 | 0.20786 | satisfactory screening result
Micro-Vickers¦SN=710-655¦scale=HV 0.05¦F=0.4905N¦mag=50X | Micro-Vickers | 710-655 | HV 0,5 | HV 0.05 | 0.05 | 0.4905 |  | 50X | 3 | 9,10,11 | 710-655_197_HMV_0.05_1_50X | 20.8492 | 0.090463 | 214.576 | 3.0205 | 20.8492 | 214.576 | 4.2145 | 1.9641 | 20.8857 | 0.14537 | 212.5323 | 2.9646 | 20.8857 | 212.5323 | 4.1871 | 1.9701 | 2.1073 | 2.0935 | -2.0437 | 5.9409 | -0.344 | 0.344 | satisfactory screening result
Micro-Vickers¦SN=711-081¦scale=HV 0.05¦F=0.4905N¦mag=50X | Micro-Vickers | 711-081 | HV 0,5 | HV 0.05 | 0.05 | 0.4905 |  | 50X | 3 | 6,7,8 | 711-081_102_HMV_0.05_1_50X | 28.7642 | 0.27597 | 112.061 | 2.1327 | 28.7642 | 112.061 | 1.9574 | 1.7467 | 28.8307 | 0.29973 | 111.551 | 2.299 | 28.8307 | 111.551 | 1.9569 | 1.7542 | 0.97869 | 0.97844 | -0.51 | 2.7678 | -0.18426 | 0.18426 | satisfactory screening result
Micro-Vickers¦SN=711-081¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers | 711-081 | HV 0,5 | HV 0.5 | 0.5 | 4.905 |  | 50X | 3 | 12,13,14 | 711-081_102_HMV_0.5_1_50X | 93.8983 | 0.68489 | 105.1503 | 1.543 | 93.8983 | 105.1503 | 1.6828 | 1.6003 | 94.0422 | 0.61296 | 104.827 | 1.3731 | 94.0422 | 104.827 | 1.6047 | 1.5308 | 0.84138 | 0.80236 | -0.32333 | 2.3253 | -0.13905 | 0.13905 | satisfactory screening result
Micro-Vickers¦SN=711-081¦scale=HV 1¦F=9.81N¦mag=20X | Micro-Vickers | 711-081 | HV 1 | HV 1 | 1 | 9.81 |  | 20X | 3 | 27,28,29 | 711-081_102_HMV_1_1_20X | 133.8632 | 0.3842 | 103.4753 | 0.58053 | 133.8632 | 103.4753 | 1.3348 | 1.29 | 133.9907 | 0.46643 | 103.2693 | 0.72108 | 133.9907 | 103.2693 | 1.3956 | 1.3514 | 0.66741 | 0.69778 | -0.206 | 1.9311 | -0.10667 | 0.10667 | satisfactory screening result
Micro-Vickers¦SN=711-081¦scale=HV 2¦F=19.62N¦mag=20X | Micro-Vickers | 711-081 | HV 2 | HV 2 | 2 | 19.62 |  | 20X | 3 | 42,43,44 | 711-081_102_HMV_2_1_20X | 189.3157 | 1.5065 | 103.472 | 1.6567 | 189.3157 | 103.472 | 2.637 | 2.5485 | 189.6583 | 1.456 | 103.097 | 1.5905 | 189.6583 | 103.097 | 2.556 | 2.4793 | 1.3185 | 1.278 | -0.375 | 3.6725 | -0.10211 | 0.10211 | satisfactory screening result
Micro-Vickers¦SN=711-246¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers | 711-246 | HV 0,5 | HV 0.5 | 0.5 | 4.905 |  | 50X | 3 | 18,19,20 | 711-246_502_HMV_0.5_1_50X | 42.603 | 0.25107 | 510.7747 | 5.9922 | 42.603 | 510.7747 | 7.7559 | 1.5185 | 42.7943 | 0.38685 | 506.258 | 9.2177 | 42.7943 | 506.258 | 8.0059 | 1.5814 | 3.878 | 4.003 | -4.5167 | 11.1467 | -0.4052 | 0.4052 | satisfactory screening result
Micro-Vickers¦SN=711-246¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers | 711-246 | HV 1 | HV 1 | 1 | 9.81 |  | 50X | 3 | 33,34,35 | 711-246_502_HMV_1_1_50X; 711-246_502_HMV_1_3_50X | 60.518 | 0.41217 | 506.2687 | 6.9094 | 60.518 | 506.2687 | 7.449 | 1.4713 | 60.277 | 0.55921 | 510.365 | 9.4016 | 60.277 | 510.365 | 8.0699 | 1.5812 | 3.7245 | 4.035 | 4.0963 | 10.9823 | 0.37299 | 0.37299 | satisfactory screening result
Micro-Vickers¦SN=711-246¦scale=HV 2¦F=19.62N¦mag=50X | Micro-Vickers | 711-246 | HV 2 | HV 2 | 2 | 19.62 |  | 50X | 3 | 48,49,50 | 711-246_502_HMV_2_1_50X | 84.9267 | 0.22055 | 514.112 | 2.676 | 84.9267 | 514.112 | 6.5828 | 1.2804 | 84.8642 | 0.26164 | 514.842 | 3.2053 | 84.8642 | 514.842 | 6.6849 | 1.2984 | 3.2914 | 3.3424 | 0.73 | 9.3819 | 0.077809 | 0.077809 | satisfactory screening result
Vickers¦SN=216-610¦scale=HV 100¦F=981N¦mag=20X | Vickers | 216-610 | HV 100 | HV 100 | 100 | 981 |  | 20X | 3 | 57,58,59 |  | 511.2327 | 0.34023 | 709.5333 | 0.94516 | 511.2327 | 709.5333 | 8.5269 | 1.2018 | 511.089 | 0.38721 | 709.9667 | 1.0693 | 511.089 | 709.9667 | 8.7636 | 1.2344 | 4.2634 | 4.3818 | 0.43333 | 12.2274 | 0.03544 | 0.03544 | satisfactory screening result
Vickers¦SN=216-610¦scale=HV 20¦F=196.2N¦mag=20X | Vickers | 216-610 | HV 20 | HV 20 | 20 | 196.2 |  | 20X | 3 | 27,28,29 | 216-610_701_HV_20_1_20X | 227.4537 | 0.39908 | 716.7313 | 2.5185 | 227.4537 | 716.7313 | 9.5727 | 1.3356 | 227.1698 | 0.48427 | 717.44 | 3.0217 | 227.1698 | 717.44 | 10.0422 | 1.3997 | 4.7863 | 5.0211 | 0.70867 | 13.8738 | 0.05108 | 0.05108 | satisfactory screening result
Vickers¦SN=216-610¦scale=HV 3¦F=29.43N¦mag=50X | Vickers | 216-610 | HV 3 | HV 3 | 3 | 29.43 |  | 50X | 3 | 12,13,14 | 216-610_701_HV_3_1_50X | 87.4893 | 0.32459 | 726.6623 | 5.3945 | 87.4893 | 726.6623 | 11.4188 | 1.5714 | 87.2083 | 0.83125 | 731.3793 | 13.7487 | 87.2083 | 731.3793 | 14.3219 | 1.9582 | 5.7094 | 7.1609 | 4.717 | 18.3168 | 0.25752 | 0.25752 | satisfactory screening result
Vickers¦SN=216-610¦scale=HV 50¦F=490.5N¦mag=20X | Vickers | 216-610 | HV 50 | HV 50 | 50 | 490.5 |  | 20X | 3 | 42,43,44 |  | 360.7707 | 0.39373 | 712.3667 | 1.5503 | 360.7707 | 712.3667 | 9.0466 | 1.2699 | 360.494 | 0.26377 | 713.4667 | 1.0599 | 360.494 | 713.4667 | 8.4852 | 1.1893 | 4.5233 | 4.2426 | 1.1 | 12.4032 | 0.088687 | 0.088687 | satisfactory screening result
Vickers¦SN=221-101¦scale=HV 100¦F=981N¦mag=20X | Vickers | 221-101 | HV 100 | HV 100 | 100 | 981 |  | 20X | 3 | 60,61,62 |  | 456.0675 | 0.16184 | 891.6 | 0.65574 | 456.0675 | 891.6 | 10.0015 | 1.1217 | 455.9177 | 0.1675 | 892.1667 | 0.65064 | 455.9177 | 892.1667 | 10.0253 | 1.1237 | 5.0007 | 5.0127 | 0.56667 | 14.1611 | 0.040016 | 0.040016 | satisfactory screening result
Vickers¦SN=221-101¦scale=HV 20¦F=196.2N¦mag=20X | Vickers | 221-101 | HV 20 | HV 20 | 20 | 196.2 |  | 20X | 3 | 30,31,32 | 221-101_898_HV_20_1_20X | 202.8217 | 0.20328 | 901.3887 | 1.8054 | 202.8217 | 901.3887 | 11.2763 | 1.251 | 202.566 | 0.37204 | 903.6693 | 3.3246 | 202.566 | 903.6693 | 12.0942 | 1.3383 | 5.6381 | 6.0471 | 2.2807 | 16.5355 | 0.13793 | 0.13793 | satisfactory screening result
Vickers¦SN=221-101¦scale=HV 3¦F=29.43N¦mag=50X | Vickers | 221-101 | HV 3 | HV 3 | 3 | 29.43 |  | 50X | 3 | 15,16,17 | 221-101_898_HV_3_1_50X | 78.1743 | 0.070671 | 910.1297 | 1.6456 | 78.1743 | 910.1297 | 14.1415 | 1.5538 | 77.9745 | 0.42186 | 914.8567 | 9.9076 | 77.9745 | 914.8567 | 15.3644 | 1.6794 | 7.0708 | 7.6822 | 4.727 | 20.8817 | 0.22637 | 0.22637 | satisfactory screening result
Vickers¦SN=221-101¦scale=HV 50¦F=490.5N¦mag=20X | Vickers | 221-101 | HV 50 | HV 50 | 50 | 490.5 |  | 20X | 3 | 45,46,47 |  | 321.8033 | 0.1286 | 895.3 | 0.72111 | 321.8033 | 895.3 | 10.2999 | 1.1504 | 321.2492 | 0.32924 | 898.3667 | 1.8475 | 321.2492 | 898.3667 | 11.1392 | 1.2399 | 5.15 | 5.5696 | 3.0667 | 15.1714 | 0.20214 | 0.20214 | satisfactory screening result
Vickers¦SN=293-250¦scale=HV 100¦F=981N¦mag=20X | Vickers | 293-250 | HV 100 | HV 100 | 100 | 981 |  | 20X | 3 | 48,49,50 |  | 1069.2362 | 3.0735 | 162.2 | 0.95394 | 1069.2362 | 162.2 | 7.7835 | 4.7987 | 1069.524 | 3.057 | 162.1333 | 0.90185 | 1069.524 | 162.1333 | 7.7406 | 4.7742 | 3.8917 | 3.8703 | -0.066667 | 10.9772 | -0.0060732 | 0.0060732 | satisfactory screening result
Vickers¦SN=293-250¦scale=HV 20¦F=196.2N¦mag=20X | Vickers | 293-250 | HV 20 | HV 20 | 20 | 196.2 |  | 20X | 3 | 18,19,20 |  | 479.2792 | 1.8649 | 161.4267 | 1.2526 | 479.2792 | 161.4267 | 4.9143 | 3.0443 | 479.3273 | 2.4103 | 161.3973 | 1.6166 | 479.3273 | 161.3973 | 6.1848 | 3.832 | 2.4571 | 3.0924 | -0.029333 | 7.8995 | -0.0037133 | 0.0037133 | satisfactory screening result
Vickers¦SN=293-250¦scale=HV 3¦F=29.43N¦mag=20X | Vickers | 293-250 | HV 3 | HV 3 | 3 | 29.43 |  | 20X | 3 | 3,4,5 | 293-250_160_HV_3_1_20X | 184.5562 | 0.61668 | 163.5923 | 0.71999 | 184.5562 | 163.5923 | 2.5363 | 1.5504 | 184.5995 | 0.20798 | 163.2407 | 0.40581 | 184.5995 | 163.2407 | 2.0781 | 1.273 | 1.2682 | 1.0391 | -0.35167 | 3.279 | -0.10725 | 0.10725 | satisfactory screening result
Vickers¦SN=293-250¦scale=HV 50¦F=490.5N¦mag=20X | Vickers | 293-250 | HV 50 | HV 50 | 50 | 490.5 |  | 20X | 3 | 33,34,35 |  | 756.2575 | 2.734 | 162.1 | 1.1533 | 756.2575 | 162.1 | 6.9689 | 4.2991 | 756.645 | 2.8734 | 162.0333 | 1.2503 | 756.645 | 162.0333 | 7.3001 | 4.5053 | 3.4844 | 3.65 | -0.066667 | 10.0924 | -0.0066056 | 0.0066056 | satisfactory screening result
Vickers¦SN=294-367¦scale=HV 100¦F=981N¦mag=20X | Vickers | 294-367 | HV 100 | HV 100 | 100 | 981 |  | 20X | 3 | 51,52,53 |  | 781.4813 | 0.26452 | 303.6667 | 0.23094 | 781.4813 | 303.6667 | 3.4297 | 1.1294 | 781.7035 | 0.24271 | 303.5 | 0.2 | 781.7035 | 303.5 | 3.3931 | 1.118 | 1.7149 | 1.6965 | -0.16667 | 4.8245 | -0.034546 | 0.034546 | satisfactory screening result
Vickers¦SN=294-367¦scale=HV 20¦F=196.2N¦mag=20X | Vickers | 294-367 | HV 20 | HV 20 | 20 | 196.2 |  | 20X | 3 | 21,22,23 |  | 349.196 | 0.35864 | 304.09 | 0.62488 | 349.196 | 304.09 | 3.7984 | 1.2491 | 349.4752 | 0.36651 | 303.6047 | 0.6368 | 349.4752 | 303.6047 | 3.8082 | 1.2543 | 1.8992 | 1.9041 | -0.48533 | 5.3787 | -0.090233 | 0.090233 | satisfactory screening result
Vickers¦SN=294-367¦scale=HV 3¦F=29.43N¦mag=20X | Vickers | 294-367 | HV 3 | HV 3 | 3 | 29.43 |  | 20X | 3 | 6,7,8 | 294-367_300_HV_3_1_20X | 135.062 | 0.2847 | 304.9003 | 1.2964 | 135.062 | 304.9003 | 4.2319 | 1.388 | 135.0908 | 0.33266 | 304.4983 | 1.5102 | 135.0908 | 304.4983 | 4.3009 | 1.4124 | 2.1159 | 2.1504 | -0.402 | 6.0338 | -0.066625 | 0.066625 | satisfactory screening result
Vickers¦SN=294-367¦scale=HV 50¦F=490.5N¦mag=20X | Vickers | 294-367 | HV 50 | HV 50 | 50 | 490.5 |  | 20X | 3 | 36,37,38 |  | 551.6933 | 0.43959 | 304.6333 | 0.47258 | 551.6933 | 304.6333 | 3.8664 | 1.2692 | 551.7327 | 0.4336 | 304.5667 | 0.46188 | 551.7327 | 304.5667 | 3.851 | 1.2644 | 1.9332 | 1.9255 | -0.066667 | 5.457 | -0.012217 | 0.012217 | satisfactory screening result
Vickers¦SN=294-481¦scale=HV 100¦F=981N¦mag=20X | Vickers | 294-481 | HV 100 | HV 100 | 100 | 981 |  | 20X | 3 | 54,55,56 |  | 605.7117 | 0.70882 | 505.4667 | 1.193 | 605.7117 | 505.4667 | 7.6895 | 1.5213 | 605.5448 | 0.92513 | 505.7333 | 1.5275 | 605.5448 | 505.7333 | 8.955 | 1.7707 | 3.8447 | 4.4775 | 0.26667 | 11.8034 | 0.022592 | 0.022592 | satisfactory screening result
Vickers¦SN=294-481¦scale=HV 20¦F=196.2N¦mag=20X | Vickers | 294-481 | HV 20 | HV 20 | 20 | 196.2 |  | 20X | 3 | 24,25,26 | 294_481_503_HV_20_1_20X | 270.0955 | 0.35338 | 508.284 | 1.3284 | 270.0955 | 508.284 | 6.497 | 1.2782 | 269.9183 | 0.2894 | 508.951 | 1.0906 | 269.9183 | 508.951 | 6.313 | 1.2404 | 3.2485 | 3.1565 | 0.667 | 9.059 | 0.073629 | 0.073629 | satisfactory screening result
Vickers¦SN=294-481¦scale=HV 3¦F=29.43N¦mag=20X | Vickers | 294-481 | HV 3 | HV 3 | 3 | 29.43 |  | 20X | 3 | 9,10,11 | 294_481_503_HV_3_1_20X | 103.873 | 0.17282 | 515.499 | 1.7126 | 103.873 | 515.499 | 7.4144 | 1.4383 | 104.2282 | 0.22259 | 511.9953 | 2.1799 | 104.2282 | 511.9953 | 7.4379 | 1.4527 | 3.7072 | 3.719 | -3.5037 | 10.5022 | -0.33361 | 0.33361 | satisfactory screening result
Vickers¦SN=294-481¦scale=HV 50¦F=490.5N¦mag=20X | Vickers | 294-481 | HV 50 | HV 50 | 50 | 490.5 |  | 20X | 3 | 39,40,41 |  | 426.5372 | 0.19549 | 509.6 | 0.45826 | 426.5372 | 509.6 | 5.8116 | 1.1404 | 426.8152 | 0.36082 | 508.9333 | 0.86217 | 426.8152 | 508.9333 | 6.2624 | 1.2305 | 2.9058 | 3.1312 | -0.66667 | 8.5435 | -0.078032 | 0.078032 | satisfactory screening result

## Repeatability
series_id | method | source | serial_number | hardness_scale | hardness_level | applied_force_n | ball_diameter_mm | magnification | n | mean_dimension_um | sd_dimension_um | cv_percent | range_dimension_um | mean_d1_minus_d2_um | mean_abs_d1_minus_d2_um
Brinell¦SN=290-736¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | Brinell | manual | 290-736 | HBW 1/30 | 1 | 294.3 | 1 | 20X | 3 | 254.9415 | 0.35075 | 0.13758 | 0.6095 | -0.26967 | 0.41033
Brinell¦SN=290-736¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | Brinell | automatic | 290-736 | HBW 1/30 | 1 | 294.3 | 1 | 20X | 3 | 255.4053 | 0.97174 | 0.38047 | 1.7085 | 0.76667 | 1.1187
Brinell¦SN=290-736¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | Brinell | manual | 290-736 | HBW 1/5 | 1 | 49.05 | 1 | 20X | 3 | 116.9922 | 1.9329 | 1.6522 | 3.4065 | -0.56767 | 0.569
Brinell¦SN=290-736¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | Brinell | automatic | 290-736 | HBW 1/5 | 1 | 49.05 | 1 | 20X | 3 | 117.9267 | 1.8837 | 1.5974 | 3.654 | 0.080667 | 0.96933
Brinell¦SN=293-250¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | Brinell | manual | 293-250 | HBW 1/30 | 1 | 294.3 | 1 | 20X | 3 | 464.74 | 0.40473 | 0.087087 | 0.785 | 0.35 | 1.636
Brinell¦SN=293-250¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | Brinell | automatic | 293-250 | HBW 1/30 | 1 | 294.3 | 1 | 20X | 3 | 466.62 | 0.9328 | 0.19991 | 1.712 | 1.138 | 2.3593
Brinell¦SN=293-250¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | Brinell | manual | 293-250 | HBW 1/5 | 1 | 49.05 | 1 | 20X | 3 | 210.4625 | 1.173 | 0.55736 | 2.1895 | -0.32367 | 0.97367
Brinell¦SN=293-250¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | Brinell | automatic | 293-250 | HBW 1/5 | 1 | 49.05 | 1 | 20X | 3 | 211.8828 | 0.70577 | 0.3331 | 1.2645 | -0.57567 | 0.819
Brinell¦SN=294-367¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | Brinell | manual | 294-367 | HBW 1/30 | 1 | 294.3 | 1 | 20X | 3 | 350.6645 | 1.1464 | 0.32692 | 2.239 | -0.44633 | 0.44633
Brinell¦SN=294-367¦scale=HBW 1/30¦F=294.3N¦ball=1mm¦mag=20X | Brinell | automatic | 294-367 | HBW 1/30 | 1 | 294.3 | 1 | 20X | 3 | 351.1102 | 1.2353 | 0.35183 | 2.271 | -0.519 | 0.613
Brinell¦SN=294-367¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | Brinell | manual | 294-367 | HBW 1/5 | 1 | 49.05 | 1 | 20X | 3 | 154.5825 | 0.18559 | 0.12006 | 0.3645 | -0.16233 | 0.649
Brinell¦SN=294-367¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | Brinell | automatic | 294-367 | HBW 1/5 | 1 | 49.05 | 1 | 20X | 3 | 155.1857 | 0.38871 | 0.25048 | 0.7595 | 0.082 | 0.398
Brinell¦SN=810-202¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | Brinell | manual | 810-202 | HBW 10/3000 | 10 | 29430 | 10 | 20X | 3 | 3804.9168 | 3.0006 | 0.07886 | 5.449 | 0.082333 | 1.2057
Brinell¦SN=810-202¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | Brinell | automatic | 810-202 | HBW 10/3000 | 10 | 29430 | 10 | 20X | 3 | 3823.7332 | 15.6461 | 0.40918 | 30.8295 | -9.2963 | 9.3777
Brinell¦SN=810-202¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=20X | Brinell | manual | 810-202 | HBW 10/500 | 10 | 4905 | 10 | 20X | 3 | 1686.2058 | 1.1722 | 0.069518 | 2.1615 | -0.040333 | 0.475
Brinell¦SN=810-202¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=20X | Brinell | automatic | 810-202 | HBW 10/500 | 10 | 4905 | 10 | 20X | 3 | 1710.5187 | 10.5711 | 0.618 | 20.7655 | -8.5747 | 22.162
Brinell¦SN=812-006¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | Brinell | manual | 812-006 | HBW 10/3000 | 10 | 29430 | 10 | 20X | 3 | 4840.7873 | 9.3405 | 0.19295 | 18.5325 | -4.8427 | 5.2807
Brinell¦SN=812-006¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | Brinell | automatic | 812-006 | HBW 10/3000 | 10 | 29430 | 10 | 20X | 3 | 4861.091 | 16.8025 | 0.34565 | 29.728 | -4.7447 | 17.8207
Brinell¦SN=812-006¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=20X | Brinell | manual | 812-006 | HBW 10/500 | 10 | 4905 | 10 | 20X | 3 | 2203.1337 | 4.495 | 0.20403 | 8.481 | 0.056 | 0.27467
Brinell¦SN=812-006¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=20X | Brinell | automatic | 812-006 | HBW 10/500 | 10 | 4905 | 10 | 20X | 3 | 2234.308 | 5.5375 | 0.24784 | 10.216 | -0.66867 | 15.6133
Brinell¦SN=8301-50DW¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | Brinell | manual | 8301-50DW | HBW 10/3000 | 10 | 29430 | 10 | 20X | 3 | 2649.1663 | 6.4182 | 0.24227 | 11.2785 | -5.0673 | 5.0673
Brinell¦SN=8301-50DW¦scale=HBW 10/3000¦F=29430N¦ball=10mm¦mag=20X | Brinell | automatic | 8301-50DW | HBW 10/3000 | 10 | 29430 | 10 | 20X | 3 | 2655.9918 | 7.5572 | 0.28453 | 15.015 | -1.8683 | 10.9143
Brinell¦SN=8301-50DW¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=10X | Brinell | manual | 8301-50DW | HBW 10/500 | 10 | 4905 | 10 | 10X | 3 | 1204.3738 | 1.3301 | 0.11044 | 2.4435 | 1.1237 | 1.1237
Brinell¦SN=8301-50DW¦scale=HBW 10/500¦F=4905N¦ball=10mm¦mag=10X | Brinell | automatic | 8301-50DW | HBW 10/500 | 10 | 4905 | 10 | 10X | 3 | 1239.9103 | 21.3949 | 1.7255 | 39.874 | 23.2767 | 40.63
Micro-Vickers¦SN=706-437¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers | manual | 706-437 | HV 0.5 | 0.5 | 4.905 |  | 50X | 3 | 53.9568 | 0.19602 | 0.36329 | 0.38 | 0.090333 | 0.090333
Micro-Vickers¦SN=706-437¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers | automatic | 706-437 | HV 0.5 | 0.5 | 4.905 |  | 50X | 3 | 54.0013 | 0.22552 | 0.41761 | 0.414 | -0.322 | 0.35133
Micro-Vickers¦SN=706-437¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers | manual | 706-437 | HV 1 | 1 | 9.81 |  | 50X | 3 | 76.676 | 1.0313 | 1.345 | 2.061 | 0.023333 | 1.0147
Micro-Vickers¦SN=706-437¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers | automatic | 706-437 | HV 1 | 1 | 9.81 |  | 50X | 3 | 76.3145 | 1.1524 | 1.5101 | 2.2975 | -0.471 | 1.2443
Micro-Vickers¦SN=706-437¦scale=HV 2¦F=19.62N¦mag=20X | Micro-Vickers | manual | 706-437 | HV 2 | 2 | 19.62 |  | 20X | 3 | 108.0538 | 0.47771 | 0.44211 | 0.9415 | -0.37367 | 0.37367
Micro-Vickers¦SN=706-437¦scale=HV 2¦F=19.62N¦mag=20X | Micro-Vickers | automatic | 706-437 | HV 2 | 2 | 19.62 |  | 20X | 3 | 108.0802 | 0.4751 | 0.43958 | 0.8975 | -0.48567 | 0.48567
Micro-Vickers¦SN=709-293¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers | manual | 709-293 | HV 0.5 | 0.5 | 4.905 |  | 50X | 3 | 30.7273 | 0.1014 | 0.33001 | 0.1905 | -0.10933 | 0.31867
Micro-Vickers¦SN=709-293¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers | automatic | 709-293 | HV 0.5 | 0.5 | 4.905 |  | 50X | 3 | 30.7503 | 0.2257 | 0.73398 | 0.434 | -0.16933 | 0.23133
Micro-Vickers¦SN=709-293¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers | manual | 709-293 | HV 1 | 1 | 9.81 |  | 50X | 3 | 45.1108 | 0.1046 | 0.23187 | 0.2085 | 0.043667 | 0.161
Micro-Vickers¦SN=709-293¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers | automatic | 709-293 | HV 1 | 1 | 9.81 |  | 50X | 3 | 44.3858 | 0.27767 | 0.62559 | 0.5305 | -0.14233 | 0.30433
Micro-Vickers¦SN=709-293¦scale=HV 2¦F=19.62N¦mag=50X | Micro-Vickers | manual | 709-293 | HV 2 | 2 | 19.62 |  | 50X | 3 | 63.2582 | 0.39976 | 0.63195 | 0.7995 | -0.19633 | 0.30433
Micro-Vickers¦SN=709-293¦scale=HV 2¦F=19.62N¦mag=50X | Micro-Vickers | automatic | 709-293 | HV 2 | 2 | 19.62 |  | 50X | 3 | 63.1747 | 0.60851 | 0.96322 | 1.2115 | -0.20133 | 0.346
Micro-Vickers¦SN=710-106¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers | manual | 710-106 | HV 0.5 | 0.5 | 4.905 |  | 50X | 3 | 35.1163 | 0.1792 | 0.5103 | 0.3155 | -0.021333 | 0.146
Micro-Vickers¦SN=710-106¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers | automatic | 710-106 | HV 0.5 | 0.5 | 4.905 |  | 50X | 3 | 35.0033 | 0.28566 | 0.81609 | 0.56 | -0.072 | 0.352
Micro-Vickers¦SN=710-106¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers | manual | 710-106 | HV 1 | 1 | 9.81 |  | 50X | 3 | 50.3497 | 0.31196 | 0.61958 | 0.6235 | -0.13133 | 0.216
Micro-Vickers¦SN=710-106¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers | automatic | 710-106 | HV 1 | 1 | 9.81 |  | 50X | 3 | 50.0255 | 0.29729 | 0.59429 | 0.592 | -0.14233 | 0.38833
Micro-Vickers¦SN=710-106¦scale=HV 2¦F=19.62N¦mag=50X | Micro-Vickers | manual | 710-106 | HV 2 | 2 | 19.62 |  | 50X | 3 | 70.8057 | 0.075836 | 0.1071 | 0.1515 | 0.0066667 | 0.28733
Micro-Vickers¦SN=710-106¦scale=HV 2¦F=19.62N¦mag=50X | Micro-Vickers | automatic | 710-106 | HV 2 | 2 | 19.62 |  | 50X | 3 | 70.5585 | 0.19166 | 0.27163 | 0.3815 | 0.24433 | 0.635
Micro-Vickers¦SN=710-404¦scale=HV 0.05¦F=0.4905N¦mag=50X | Micro-Vickers | manual | 710-404 | HV 0.05 | 0.05 | 0.4905 |  | 50X | 3 | 43.3358 | 1.0385 | 2.3964 | 2.025 | 0.80833 | 0.80833
Micro-Vickers¦SN=710-404¦scale=HV 0.05¦F=0.4905N¦mag=50X | Micro-Vickers | automatic | 710-404 | HV 0.05 | 0.05 | 0.4905 |  | 50X | 3 | 43.4723 | 1.0312 | 2.3721 | 2.028 | 0.618 | 0.618
Micro-Vickers¦SN=710-655¦scale=HV 0.05¦F=0.4905N¦mag=50X | Micro-Vickers | manual | 710-655 | HV 0.05 | 0.05 | 0.4905 |  | 50X | 3 | 20.8492 | 0.090463 | 0.43389 | 0.1805 | -0.24967 | 0.311
Micro-Vickers¦SN=710-655¦scale=HV 0.05¦F=0.4905N¦mag=50X | Micro-Vickers | automatic | 710-655 | HV 0.05 | 0.05 | 0.4905 |  | 50X | 3 | 20.8857 | 0.14537 | 0.69603 | 0.289 | -0.35533 | 0.378
Micro-Vickers¦SN=711-081¦scale=HV 0.05¦F=0.4905N¦mag=50X | Micro-Vickers | manual | 711-081 | HV 0.05 | 0.05 | 0.4905 |  | 50X | 3 | 28.7642 | 0.27597 | 0.95942 | 0.49 | 0.471 | 0.471
Micro-Vickers¦SN=711-081¦scale=HV 0.05¦F=0.4905N¦mag=50X | Micro-Vickers | automatic | 711-081 | HV 0.05 | 0.05 | 0.4905 |  | 50X | 3 | 28.8307 | 0.29973 | 1.0396 | 0.55 | 0.59267 | 0.59267
Micro-Vickers¦SN=711-081¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers | manual | 711-081 | HV 0.5 | 0.5 | 4.905 |  | 50X | 3 | 93.8983 | 0.68489 | 0.72939 | 1.2615 | 0.22067 | 0.29467
Micro-Vickers¦SN=711-081¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers | automatic | 711-081 | HV 0.5 | 0.5 | 4.905 |  | 50X | 3 | 94.0422 | 0.61296 | 0.65179 | 1.1745 | 0.373 | 0.373
Micro-Vickers¦SN=711-081¦scale=HV 1¦F=9.81N¦mag=20X | Micro-Vickers | manual | 711-081 | HV 1 | 1 | 9.81 |  | 20X | 3 | 133.8632 | 0.3842 | 0.28701 | 0.7625 | 0.045 | 0.63833
Micro-Vickers¦SN=711-081¦scale=HV 1¦F=9.81N¦mag=20X | Micro-Vickers | automatic | 711-081 | HV 1 | 1 | 9.81 |  | 20X | 3 | 133.9907 | 0.46643 | 0.34811 | 0.868 | 0.065333 | 0.65467
Micro-Vickers¦SN=711-081¦scale=HV 2¦F=19.62N¦mag=20X | Micro-Vickers | manual | 711-081 | HV 2 | 2 | 19.62 |  | 20X | 3 | 189.3157 | 1.5065 | 0.79574 | 2.8105 | 1.7067 | 1.7067
Micro-Vickers¦SN=711-081¦scale=HV 2¦F=19.62N¦mag=20X | Micro-Vickers | automatic | 711-081 | HV 2 | 2 | 19.62 |  | 20X | 3 | 189.6583 | 1.456 | 0.76767 | 2.7875 | 1.5853 | 1.5853
Micro-Vickers¦SN=711-246¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers | manual | 711-246 | HV 0.5 | 0.5 | 4.905 |  | 50X | 3 | 42.603 | 0.25107 | 0.58933 | 0.466 | 0.060667 | 0.10733
Micro-Vickers¦SN=711-246¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers | automatic | 711-246 | HV 0.5 | 0.5 | 4.905 |  | 50X | 3 | 42.7943 | 0.38685 | 0.90397 | 0.7025 | -0.14133 | 0.31933
Micro-Vickers¦SN=711-246¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers | manual | 711-246 | HV 1 | 1 | 9.81 |  | 50X | 3 | 60.518 | 0.41217 | 0.68107 | 0.8195 | 0.512 | 0.76467
Micro-Vickers¦SN=711-246¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers | automatic | 711-246 | HV 1 | 1 | 9.81 |  | 50X | 3 | 60.277 | 0.55921 | 0.92774 | 0.983 | 0.36 | 0.61533
Micro-Vickers¦SN=711-246¦scale=HV 2¦F=19.62N¦mag=50X | Micro-Vickers | manual | 711-246 | HV 2 | 2 | 19.62 |  | 50X | 3 | 84.9267 | 0.22055 | 0.25969 | 0.382 | -0.08 | 0.14667
Micro-Vickers¦SN=711-246¦scale=HV 2¦F=19.62N¦mag=50X | Micro-Vickers | automatic | 711-246 | HV 2 | 2 | 19.62 |  | 50X | 3 | 84.8642 | 0.26164 | 0.30831 | 0.4725 | -0.14367 | 0.20367
Vickers¦SN=216-610¦scale=HV 100¦F=981N¦mag=20X | Vickers | manual | 216-610 | HV 100 | 100 | 981 |  | 20X | 3 | 511.2327 | 0.34023 | 0.06655 | 0.651 | -0.20467 | 0.302
Vickers¦SN=216-610¦scale=HV 100¦F=981N¦mag=20X | Vickers | automatic | 216-610 | HV 100 | 100 | 981 |  | 20X | 3 | 511.089 | 0.38721 | 0.075762 | 0.69 | 0.146 | 0.146
Vickers¦SN=216-610¦scale=HV 20¦F=196.2N¦mag=20X | Vickers | manual | 216-610 | HV 20 | 20 | 196.2 |  | 20X | 3 | 227.4537 | 0.39908 | 0.17545 | 0.7265 | 0.42333 | 0.42333
Vickers¦SN=216-610¦scale=HV 20¦F=196.2N¦mag=20X | Vickers | automatic | 216-610 | HV 20 | 20 | 196.2 |  | 20X | 3 | 227.1698 | 0.48427 | 0.21317 | 0.915 | -0.37233 | 0.48767
Vickers¦SN=216-610¦scale=HV 3¦F=29.43N¦mag=50X | Vickers | manual | 216-610 | HV 3 | 3 | 29.43 |  | 50X | 3 | 87.4893 | 0.32459 | 0.37101 | 0.6485 | -0.19867 | 0.23867
Vickers¦SN=216-610¦scale=HV 3¦F=29.43N¦mag=50X | Vickers | automatic | 216-610 | HV 3 | 3 | 29.43 |  | 50X | 3 | 87.2083 | 0.83125 | 0.95318 | 1.5305 | -0.35067 | 0.35067
Vickers¦SN=216-610¦scale=HV 50¦F=490.5N¦mag=20X | Vickers | manual | 216-610 | HV 50 | 50 | 490.5 |  | 20X | 3 | 360.7707 | 0.39373 | 0.10914 | 0.787 | -0.018667 | 0.12933
Vickers¦SN=216-610¦scale=HV 50¦F=490.5N¦mag=20X | Vickers | automatic | 216-610 | HV 50 | 50 | 490.5 |  | 20X | 3 | 360.494 | 0.26377 | 0.073168 | 0.522 | 0.486 | 0.486
Vickers¦SN=221-101¦scale=HV 100¦F=981N¦mag=20X | Vickers | manual | 221-101 | HV 100 | 100 | 981 |  | 20X | 3 | 456.0675 | 0.16184 | 0.035486 | 0.321 | -0.24367 | 0.24367
Vickers¦SN=221-101¦scale=HV 100¦F=981N¦mag=20X | Vickers | automatic | 221-101 | HV 100 | 100 | 981 |  | 20X | 3 | 455.9177 | 0.1675 | 0.03674 | 0.335 | -0.148 | 0.494
Vickers¦SN=221-101¦scale=HV 20¦F=196.2N¦mag=20X | Vickers | manual | 221-101 | HV 20 | 20 | 196.2 |  | 20X | 3 | 202.8217 | 0.20328 | 0.10023 | 0.3685 | 0.37267 | 0.37267
Vickers¦SN=221-101¦scale=HV 20¦F=196.2N¦mag=20X | Vickers | automatic | 221-101 | HV 20 | 20 | 196.2 |  | 20X | 3 | 202.566 | 0.37204 | 0.18366 | 0.663 | 0.033333 | 0.58667
Vickers¦SN=221-101¦scale=HV 3¦F=29.43N¦mag=50X | Vickers | manual | 221-101 | HV 3 | 3 | 29.43 |  | 50X | 3 | 78.1743 | 0.070671 | 0.090401 | 0.141 | -0.046667 | 0.046667
Vickers¦SN=221-101¦scale=HV 3¦F=29.43N¦mag=50X | Vickers | automatic | 221-101 | HV 3 | 3 | 29.43 |  | 50X | 3 | 77.9745 | 0.42186 | 0.54102 | 0.842 | -0.14233 | 0.35633
Vickers¦SN=221-101¦scale=HV 50¦F=490.5N¦mag=20X | Vickers | manual | 221-101 | HV 50 | 50 | 490.5 |  | 20X | 3 | 321.8033 | 0.1286 | 0.039962 | 0.246 | 0.11667 | 0.24467
Vickers¦SN=221-101¦scale=HV 50¦F=490.5N¦mag=20X | Vickers | automatic | 221-101 | HV 50 | 50 | 490.5 |  | 20X | 3 | 321.2492 | 0.32924 | 0.10249 | 0.5715 | 0.40767 | 0.40767
Vickers¦SN=293-250¦scale=HV 100¦F=981N¦mag=20X | Vickers | manual | 293-250 | HV 100 | 100 | 981 |  | 20X | 3 | 1069.2362 | 3.0735 | 0.28745 | 6.126 | -2.6097 | 2.6097
Vickers¦SN=293-250¦scale=HV 100¦F=981N¦mag=20X | Vickers | automatic | 293-250 | HV 100 | 100 | 981 |  | 20X | 3 | 1069.524 | 3.057 | 0.28583 | 6.0985 | -2.4487 | 2.55
Vickers¦SN=293-250¦scale=HV 20¦F=196.2N¦mag=20X | Vickers | manual | 293-250 | HV 20 | 20 | 196.2 |  | 20X | 3 | 479.2792 | 1.8649 | 0.38912 | 3.431 | 0.193 | 0.57833
Vickers¦SN=293-250¦scale=HV 20¦F=196.2N¦mag=20X | Vickers | automatic | 293-250 | HV 20 | 20 | 196.2 |  | 20X | 3 | 479.3273 | 2.4103 | 0.50285 | 4.3805 | 0.29067 | 0.79667
Vickers¦SN=293-250¦scale=HV 3¦F=29.43N¦mag=20X | Vickers | manual | 293-250 | HV 3 | 3 | 29.43 |  | 20X | 3 | 184.5562 | 0.61668 | 0.33414 | 1.232 | -0.64433 | 0.67033
Vickers¦SN=293-250¦scale=HV 3¦F=29.43N¦mag=20X | Vickers | automatic | 293-250 | HV 3 | 3 | 29.43 |  | 20X | 3 | 184.5995 | 0.20798 | 0.11267 | 0.3675 | -0.30167 | 0.31367
Vickers¦SN=293-250¦scale=HV 50¦F=490.5N¦mag=20X | Vickers | manual | 293-250 | HV 50 | 50 | 490.5 |  | 20X | 3 | 756.2575 | 2.734 | 0.36152 | 5.4555 | -0.32967 | 0.511
Vickers¦SN=293-250¦scale=HV 50¦F=490.5N¦mag=20X | Vickers | automatic | 293-250 | HV 50 | 50 | 490.5 |  | 20X | 3 | 756.645 | 2.8734 | 0.37975 | 5.6995 | -0.4 | 0.4
Vickers¦SN=294-367¦scale=HV 100¦F=981N¦mag=20X | Vickers | manual | 294-367 | HV 100 | 100 | 981 |  | 20X | 3 | 781.4813 | 0.26452 | 0.033849 | 0.469 | -1.3893 | 1.3893
Vickers¦SN=294-367¦scale=HV 100¦F=981N¦mag=20X | Vickers | automatic | 294-367 | HV 100 | 100 | 981 |  | 20X | 3 | 781.7035 | 0.24271 | 0.031048 | 0.4815 | -0.92433 | 0.92433
Vickers¦SN=294-367¦scale=HV 20¦F=196.2N¦mag=20X | Vickers | manual | 294-367 | HV 20 | 20 | 196.2 |  | 20X | 3 | 349.196 | 0.35864 | 0.10271 | 0.7135 | -0.036 | 0.13533
Vickers¦SN=294-367¦scale=HV 20¦F=196.2N¦mag=20X | Vickers | automatic | 294-367 | HV 20 | 20 | 196.2 |  | 20X | 3 | 349.4752 | 0.36651 | 0.10488 | 0.7185 | -0.019667 | 0.41033
Vickers¦SN=294-367¦scale=HV 3¦F=29.43N¦mag=20X | Vickers | manual | 294-367 | HV 3 | 3 | 29.43 |  | 20X | 3 | 135.062 | 0.2847 | 0.21079 | 0.568 | 0.372 | 0.372
Vickers¦SN=294-367¦scale=HV 3¦F=29.43N¦mag=20X | Vickers | automatic | 294-367 | HV 3 | 3 | 29.43 |  | 20X | 3 | 135.0908 | 0.33266 | 0.24625 | 0.6585 | 0.351 | 0.599
Vickers¦SN=294-367¦scale=HV 50¦F=490.5N¦mag=20X | Vickers | manual | 294-367 | HV 50 | 50 | 490.5 |  | 20X | 3 | 551.6933 | 0.43959 | 0.07968 | 0.8215 | -0.754 | 0.78
Vickers¦SN=294-367¦scale=HV 50¦F=490.5N¦mag=20X | Vickers | automatic | 294-367 | HV 50 | 50 | 490.5 |  | 20X | 3 | 551.7327 | 0.4336 | 0.078589 | 0.754 | -0.59667 | 1.0673
Vickers¦SN=294-481¦scale=HV 100¦F=981N¦mag=20X | Vickers | manual | 294-481 | HV 100 | 100 | 981 |  | 20X | 3 | 605.7117 | 0.70882 | 0.11702 | 1.3665 | -0.834 | 0.834
Vickers¦SN=294-481¦scale=HV 100¦F=981N¦mag=20X | Vickers | automatic | 294-481 | HV 100 | 100 | 981 |  | 20X | 3 | 605.5448 | 0.92513 | 0.15278 | 1.823 | -0.82967 | 0.82967
Vickers¦SN=294-481¦scale=HV 20¦F=196.2N¦mag=20X | Vickers | manual | 294-481 | HV 20 | 20 | 196.2 |  | 20X | 3 | 270.0955 | 0.35338 | 0.13084 | 0.6405 | 0.16767 | 0.19833
Vickers¦SN=294-481¦scale=HV 20¦F=196.2N¦mag=20X | Vickers | automatic | 294-481 | HV 20 | 20 | 196.2 |  | 20X | 3 | 269.9183 | 0.2894 | 0.10722 | 0.5785 | 0.15133 | 0.268
Vickers¦SN=294-481¦scale=HV 3¦F=29.43N¦mag=20X | Vickers | manual | 294-481 | HV 3 | 3 | 29.43 |  | 20X | 3 | 103.873 | 0.17282 | 0.16637 | 0.3185 | -0.15933 | 0.376
Vickers¦SN=294-481¦scale=HV 3¦F=29.43N¦mag=20X | Vickers | automatic | 294-481 | HV 3 | 3 | 29.43 |  | 20X | 3 | 104.2282 | 0.22259 | 0.21356 | 0.4085 | -0.039 | 0.333
Vickers¦SN=294-481¦scale=HV 50¦F=490.5N¦mag=20X | Vickers | manual | 294-481 | HV 50 | 50 | 490.5 |  | 20X | 3 | 426.5372 | 0.19549 | 0.045832 | 0.385 | -0.349 | 0.447
Vickers¦SN=294-481¦scale=HV 50¦F=490.5N¦mag=20X | Vickers | automatic | 294-481 | HV 50 | 50 | 490.5 |  | 20X | 3 | 426.8152 | 0.36082 | 0.084539 | 0.7085 | 0.049667 | 0.683

## Repeatability_Summary
method | source | magnification | series_count | mean_sd_um | median_sd_um | mean_abs_d1_minus_d2_um
Brinell | automatic | 10X | 1 | 21.3949 | 21.3949 | 40.63
Brinell | automatic | 20X | 11 | 5.6575 | 1.8837 | 7.4696
Brinell | manual | 10X | 1 | 1.3301 | 1.3301 | 1.1237
Brinell | manual | 20X | 11 | 2.6927 | 1.173 | 1.5443
Micro-Vickers | automatic | 20X | 3 | 0.79916 | 0.4751 | 0.90856
Micro-Vickers | automatic | 50X | 15 | 0.43743 | 0.29729 | 0.46351
Micro-Vickers | manual | 20X | 3 | 0.78946 | 0.47771 | 0.90622
Micro-Vickers | manual | 50X | 15 | 0.35824 | 0.25107 | 0.3628
Vickers | automatic | 20X | 18 | 0.76256 | 0.36367 | 0.65461
Vickers | automatic | 50X | 2 | 0.62655 | 0.62655 | 0.3535
Vickers | manual | 20X | 18 | 0.70521 | 0.35601 | 0.58983
Vickers | manual | 50X | 2 | 0.19763 | 0.19763 | 0.14267

## Calibration
image_path | magnification | orientation | detected_step_um | um_per_pixel | intercept_um | n_ticks | x_start_px | x_end_px | detected_span_px | detected_span_um | residual_sd_um | residual_max_abs_um | full_scale_predicted_px | full_scale_visible | full_scale_consistency_percent | status | reason
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_20x.bmp | 20X | horizontal | 10 | 0.2439 | -72.6829 | 11 | 298 | 708 | 410 | 100 | 0 | 0 | 410 | TRUE | 100 | reliable | 11 ticks fitted as 10 um lattice using fixed_ruler_band; span covers 100.0% of configured 100 um scale
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | horizontal | 2 | 0.097649 | -1.9568 | 51 | 20 | 1044 | 1024 | 100 | 0.03215 | 0.076501 | 1024.0781 | TRUE | 100 | reliable | 51 ticks fitted as 2 um lattice using fixed_ruler_band; span covers 100.0% of configured 100 um scale

## Calibration_Ticks
image_path | magnification | tick_index | pixel_position | assigned_position_um | fit_residual_um
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_20x.bmp | 20X | 0 | 298 | 0 | 0
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_20x.bmp | 20X | 1 | 339 | 10 | 0
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_20x.bmp | 20X | 2 | 380 | 20 | 0
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_20x.bmp | 20X | 3 | 421 | 30 | 0
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_20x.bmp | 20X | 4 | 462 | 40 | 0
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_20x.bmp | 20X | 5 | 503 | 50 | 0
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_20x.bmp | 20X | 6 | 544 | 60 | 0
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_20x.bmp | 20X | 7 | 585 | 70 | 0
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_20x.bmp | 20X | 8 | 626 | 80 | 0
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_20x.bmp | 20X | 9 | 667 | 90 | 0
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_20x.bmp | 20X | 10 | 708 | 100 | 0
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 0 | 20 | 0 | 0.0038445
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 1 | 40 | 2 | 0.050869
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 2 | 61 | 4 | 0.00024372
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 3 | 81 | 6 | 0.047268
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 4 | 102 | 8 | -0.0033571
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 5 | 122 | 10 | 0.043667
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 6 | 143 | 12 | -0.0069579
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 7 | 163 | 14 | 0.040066
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 8 | 184 | 16 | -0.010559
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 9 | 205 | 18 | -0.061184
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 10 | 225 | 20 | -0.01416
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 11 | 245 | 22 | 0.032864
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 12 | 266 | 24 | -0.01776
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 13 | 287 | 26 | -0.068385
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 14 | 307 | 28 | -0.021361
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 15 | 327 | 30 | 0.025663
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 16 | 348 | 32 | -0.024962
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 17 | 368 | 34 | 0.022062
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 18 | 389 | 36 | -0.028563
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 19 | 409 | 38 | 0.018461
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 20 | 430 | 40 | -0.032164
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 21 | 450 | 42 | 0.01486
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 22 | 471 | 44 | -0.035764
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 23 | 491 | 46 | 0.01126
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 24 | 512 | 48 | -0.039365
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 25 | 532 | 50 | 0.0076587
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 26 | 553 | 52 | -0.042966
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 27 | 573 | 54 | 0.0040579
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 28 | 594 | 56 | -0.046567
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 29 | 614 | 58 | 0.0004571
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 30 | 635 | 60 | -0.050168
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 31 | 655 | 62 | -0.0031437
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 32 | 675 | 64 | 0.04388
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 33 | 696 | 66 | -0.0067445
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 34 | 716 | 68 | 0.040279
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 35 | 737 | 70 | -0.010345
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 36 | 757 | 72 | 0.036679
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 37 | 778 | 74 | -0.013946
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 38 | 798 | 76 | 0.033078
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 39 | 819 | 78 | -0.017547
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 40 | 839 | 80 | 0.029477
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 41 | 859 | 82 | 0.076501
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 42 | 880 | 84 | 0.025876
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 43 | 901 | 86 | -0.024749
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 44 | 921 | 88 | 0.022275
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 45 | 942 | 90 | -0.028349
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 46 | 962 | 92 | 0.018675
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 47 | 983 | 94 | -0.03195
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 48 | 1003 | 96 | 0.015074
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 49 | 1024 | 98 | -0.035551
C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\calibration\stage_micrometer_0.002mm_50x.bmp | 50X | 50 | 1044 | 100 | 0.011473

## Image_Measurements
record_id | series_id | method | picture_id | excel_row | magnification | manual_mean_dimension_um | automatic_mean_dimension_um | manual_hardness | automatic_hardness | image_path | image_status | classification_reason | analysis_type | contour_area_px2 | contour_perimeter_px | bbox_x | bbox_y | bbox_w | bbox_h | bbox_aspect | contour_extent | circularity | centrality | corner_count | corner_radius_cv | min_corner_angle_deg | max_corner_angle_deg | edge_strength | python_d1_um | python_d2_um | python_mean_dimension_um | diagonal_ratio | corner_points | python_hardness | python_minus_manual_dimension_um | python_minus_automatic_dimension_um | python_minus_manual_relative_percent | python_minus_automatic_relative_percent | circle_radius_um | radial_residual_sd_um | radial_cv | ellipse_major_um | ellipse_minor_um | ellipse_axis_ratio | diameter_direction_range_4_um | diameter_direction_range_6_um | diameter_0deg_um | diameter_45deg_um | diameter_90deg_um | diameter_135deg_um | diameter6_0deg_um | diameter6_30deg_um | diameter6_60deg_um | diameter6_90deg_um | diameter6_120deg_um | diameter6_150deg_um
Vickers¦293-250¦HV 3¦1¦row3 | Vickers¦SN=293-250¦scale=HV 3¦F=29.43N¦mag=20X | Vickers | 293-250_160_HV_3_1_20X | 3 | 20X | 183.9235 | 184.3595 | 164.421 | 163.709 | C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\indentations\293-250_160_HV_3_1_20X.bmp | conditionally reliable | all configured image checks passed; elevated agreement difference | Vickers contour-corner | 336063 | 5800.0714 | 124 | 66 | 786 | 793 | 1.0089 | 0.53917 | 0.12553 | 0.97372 | 4 | 0.086675 | 48.0967 | 166.0368 | 132.8345 | 189.7382 | 192.454 | 191.0961 | 1.0143 | [[124.0, 419.0], [484.0, 71.0], [793.0, 816.0], [427.0, 858.0]] | 152.3976 | 7.1726 | 6.7366 | 3.8252 | 3.5885
Vickers¦294-367¦HV 3¦1¦row6 | Vickers¦SN=294-367¦scale=HV 3¦F=29.43N¦mag=20X | Vickers | 294-367_300_HV_3_1_20X | 6 | 20X | 135.3575 | 135.4475 | 303.553 | 303.17 | C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\indentations\294-367_300_HV_3_1_20X.bmp | rejected | all configured image checks passed; excessive agreement difference | Vickers contour-corner | 315198.5 | 5360.5454 | 205 | 135 | 683 | 658 | 1.038 | 0.70135 | 0.13784 | 0.96468 | 4 | 0.058921 | 50.1744 | 128.1234 | 61.6117 | 165.5537 | 194.2451 | 179.8994 | 1.1733 | [[335.0, 157.0], [839.0, 244.0], [887.0, 552.0], [224.0, 750.0]] | 171.9579 | 44.5419 | 44.4519 | 28.2575 | 28.1924
Vickers¦294-481¦HV 3¦1¦row9 | Vickers¦SN=294-481¦scale=HV 3¦F=29.43N¦mag=20X | Vickers | 294_481_503_HV_3_1_20X | 9 | 20X | 103.7525 | 104.126 | 516.694 | 512.994 | C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\indentations\294_481_503_HV_3_1_20X.bmp | rejected | all configured image checks passed; excessive agreement difference | Vickers contour-corner | 250436 | 4921.6795 | 220 | 185 | 657 | 578 | 1.1367 | 0.65948 | 0.12992 | 0.95218 | 4 | 0.066701 | 67.0558 | 117.7324 | 22.6716 | 165.302 | 170.3803 | 167.8411 | 1.0307 | [[286.0, 248.0], [864.0, 335.0], [791.0, 700.0], [241.0, 651.0]] | 197.5536 | 64.0886 | 63.7151 | 47.1945 | 46.855
Vickers¦216-610¦HV 3¦1¦row12 | Vickers¦SN=216-610¦scale=HV 3¦F=29.43N¦mag=50X | Vickers | 216-610_701_HV_3_1_50X | 12 | 50X | 87.1565 | 86.6305 | 732.203 | 740.873 | C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\indentations\216-610_701_HV_3_1_50X.bmp | rejected | all configured image checks passed; excessive agreement difference | Vickers contour-corner | 2650 | 347.7401 | 275 | 66 | 79 | 68 | 1.1618 | 0.4933 | 0.27539 | 0.4104 | 4 | 0.13693 | 51.9658 | 153.6587 | 90.3264 | 7.6566 | 7.6197 | 7.6381 | 1.0048 | [[275.0, 81.0], [308.0, 66.0], [353.0, 89.0], [348.0, 133.0]] | 95390.5879 | -79.5184 | -78.9924 | -167.7697 | -167.5899
Vickers¦221-101¦HV 3¦1¦row15 | Vickers¦SN=221-101¦scale=HV 3¦F=29.43N¦mag=50X | Vickers | 221-101_898_HV_3_1_50X | 15 | 50X | 78.242 | 78.0055 | 908.555 | 914.084 | C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\indentations\221-101_898_HV_3_1_50X.bmp | reliable | all configured image checks passed | Vickers contour-corner | 331171 | 3065.513 | 194 | 46 | 791 | 783 | 1.0102 | 0.5347 | 0.44285 | 0.9099 | 4 | 0.055869 | 80.9002 | 97.4423 | 26.0768 | 79.5297 | 79.5763 | 79.553 | 1.0006 | [[493.0, 46.0], [984.0, 356.0], [724.0, 827.0], [194.0, 556.0]] | 879.3646 | 1.311 | 1.5475 | 1.6616 | 1.9643
Vickers¦294-481¦HV 20¦1¦row24 | Vickers¦SN=294-481¦scale=HV 20¦F=196.2N¦mag=20X | Vickers | 294_481_503_HV_20_1_20X | 24 | 20X | 269.923 | 269.9075 | 508.932 | 508.993 | C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\indentations\294_481_503_HV_20_1_20X.bmp | rejected | No suitable indentation contour detected
Vickers¦216-610¦HV 20¦1¦row27 | Vickers¦SN=216-610¦scale=HV 20¦F=196.2N¦mag=20X | Vickers | 216-610_701_HV_20_1_20X | 27 | 20X | 227.7215 | 227.719 | 715.042 | 715.058 | C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\indentations\216-610_701_HV_20_1_20X.bmp | reliable | all configured image checks passed | Vickers contour-corner | 428740 | 3032.9099 | 101 | 24 | 847 | 843 | 1.0047 | 0.60046 | 0.58571 | 0.99379 | 4 | 0.057734 | 82.0894 | 98.449 | 171.7323 | 227.4663 | 228.2432 | 227.8547 | 1.0034 | [[333.0, 24.0], [947.0, 252.0], [734.0, 866.0], [101.0, 652.0]] | 714.6187 | 0.13325 | 0.13575 | 0.058495 | 0.059593
Vickers¦221-101¦HV 20¦1¦row30 | Vickers¦SN=221-101¦scale=HV 20¦F=196.2N¦mag=20X | Vickers | 221-101_898_HV_20_1_20X | 30 | 20X | 202.687 | 202.761 | 902.585 | 901.926 | C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\indentations\221-101_898_HV_20_1_20X.bmp | conditionally reliable | all configured image checks passed; elevated agreement difference | Vickers contour-corner | 361943.5 | 3026.9962 | 105 | 25 | 841 | 839 | 1.0024 | 0.51296 | 0.49639 | 0.99379 | 4 | 0.044083 | 82.2368 | 94.3597 | 17.0294 | 207.8571 | 207.1701 | 207.5136 | 1.0033 | [[450.0, 25.0], [945.0, 388.0], [605.0, 863.0], [105.0, 514.0]] | 861.5835 | 4.8266 | 4.7526 | 2.3533 | 2.3168
Micro-Vickers¦710-404¦HV 0.05¦1¦row3 | Micro-Vickers¦SN=710-404¦scale=HV 0.05¦F=0.4905N¦mag=50X | Micro-Vickers | 710-404_39.6_HMV_0.05_1_50X | 3 | 50X | 44.215 | 44.378 | 47.418 | 47.07 | C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\indentations\710-404_39.6_HMV_0.05_1_50X.bmp | rejected | all configured image checks passed; excessive agreement difference | Vickers contour-corner | 95839 | 2080.5525 | 271 | 215 | 485 | 580 | 1.1959 | 0.3407 | 0.27822 | 0.91506 | 4 | 0.27271 | 75.7637 | 104.0971 | 96.7884 | 59.2369 | 47.6119 | 53.4244 | 1.2442 | [[474.0, 215.0], [755.0, 370.0], [655.0, 794.0], [271.0, 429.0]] | 32.4976 | 9.2094 | 9.0464 | 18.864 | 18.4993
Micro-Vickers¦711-081¦HV 0.05¦1¦row6 | Micro-Vickers¦SN=711-081¦scale=HV 0.05¦F=0.4905N¦mag=50X | Micro-Vickers | 711-081_102_HMV_0.05_1_50X | 6 | 50X | 28.5925 | 28.6245 | 113.39 | 113.141 | C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\indentations\711-081_102_HMV_0.05_1_50X.bmp | rejected | all configured image checks passed; excessive agreement difference | Vickers contour-corner | 112042 | 4495.1555 | 255 | 220 | 568 | 473 | 1.2008 | 0.41703 | 0.069679 | 0.97992 | 4 | 0.20764 | 58.1812 | 159.2406 | 55.118 | 48.9268 | 51.7506 | 50.3387 | 1.0577 | [[255.0, 384.0], [822.0, 330.0], [734.0, 531.0], [434.0, 691.0]] | 36.6038 | 21.7462 | 21.7142 | 55.1017 | 54.9983
Micro-Vickers¦710-655¦HV 0.05¦1¦row9 | Micro-Vickers¦SN=710-655¦scale=HV 0.05¦F=0.4905N¦mag=50X | Micro-Vickers | 710-655_197_HMV_0.05_1_50X | 9 | 50X | 20.842 | 20.904 | 217.337 | 212.139 | C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\indentations\710-655_197_HMV_0.05_1_50X.bmp | rejected | all configured image checks passed; excessive agreement difference | Vickers contour-corner | 30343 | 1294.8082 | 390 | 274 | 246 | 302 | 1.2276 | 0.40843 | 0.22744 | 0.96035 | 4 | 0.18613 | 47.5464 | 161.2828 | 100.0649 | 27.9231 | 26.3053 | 27.1142 | 1.0615 | [[571.0, 274.0], [635.0, 463.0], [518.0, 555.0], [390.0, 575.0]] | 126.1646 | 6.2722 | 6.2102 | 26.158 | 25.866
Micro-Vickers¦711-081¦HV 0.5¦1¦row12 | Micro-Vickers¦SN=711-081¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers | 711-081_102_HMV_0.5_1_50X | 12 | 50X | 94.375 | 94.528 | 104.08 | 103.743 | C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\indentations\711-081_102_HMV_0.5_1_50X.bmp | rejected | No suitable indentation contour detected
Micro-Vickers¦706-437¦HV 0.5¦1¦row15 | Micro-Vickers¦SN=706-437¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers | 706-437_307_HMV_0.5_1_50X | 15 | 50X | 54.0125 | 53.898 | 318.225 | 319.106 | C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\indentations\706-437_307_HMV_0.5_1_50X.bmp | conditionally reliable | all configured image checks passed; elevated agreement difference | Vickers contour-corner | 170295 | 3008.4591 | 253 | 141 | 552 | 553 | 1.0018 | 0.55788 | 0.23644 | 0.95593 | 4 | 0.062965 | 54.6586 | 118.7742 | 105.8017 | 51.0923 | 54.3139 | 52.7031 | 1.0631 | [[671.0, 164.0], [804.0, 515.0], [613.0, 684.0], [253.0, 439.0]] | 333.9325 | -1.3094 | -1.1949 | -2.454 | -2.2419
Micro-Vickers¦711-246¦HV 0.5¦1¦row18 | Micro-Vickers¦SN=711-246¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers | 711-246_502_HMV_0.5_1_50X | 18 | 50X | 42.89 | 43.052 | 503.927 | 500.142 | C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\indentations\711-246_502_HMV_0.5_1_50X.bmp | rejected | all configured image checks passed; excessive agreement difference | Vickers contour-corner | 217562 | 7052.2421 | 216 | 143 | 686 | 613 | 1.1191 | 0.51737 | 0.054972 | 0.95518 | 4 | 0.15024 | 67.9303 | 106.3691 | 29.2913 | 64.6073 | 71.2497 | 67.9285 | 1.1028 | [[295.0, 256.0], [806.0, 146.0], [828.0, 648.0], [234.0, 599.0]] | 201.0141 | 25.0385 | 24.8765 | 45.1883 | 44.8304
Micro-Vickers¦710-106¦HV 0.5¦1¦row21 | Micro-Vickers¦SN=710-106¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers | 710-106_708_HMV_0.5_1_50X | 21 | 50X | 35.2145 | 34.938 | 747.544 | 759.423 | C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\indentations\710-106_708_HMV_0.5_1_50X.bmp | rejected | all configured image checks passed; excessive agreement difference | Vickers contour-corner | 79888 | 1623.5778 | 350 | 257 | 381 | 376 | 1.0133 | 0.55766 | 0.38084 | 0.98144 | 4 | 0.046667 | 45.0231 | 171.8078 | 30.9515 | 29.8454 | 33.6333 | 31.7394 | 1.1269 | [[727.0, 378.0], [683.0, 553.0], [557.0, 632.0], [350.0, 465.0]] | 920.7339 | -3.4751 | -3.1986 | -10.3807 | -9.5944
Micro-Vickers¦709-293¦HV 0.5¦1¦row24 | Micro-Vickers¦SN=709-293¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers | 709-293_904_HMV_0.5_1_50X | 24 | 50X | 30.8025 | 30.9315 | 977.03 | 968.929 | C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\indentations\709-293_904_HMV_0.5_1_50X.bmp | rejected | all configured image checks passed; excessive agreement difference | Vickers contour-corner | 60891 | 1144.5971 | 359 | 275 | 331 | 355 | 1.0725 | 0.5182 | 0.58406 | 0.99177 | 4 | 0.064842 | 77.1666 | 113.6293 | 35.9303 | 34.6406 | 32.2336 | 33.4371 | 1.0747 | [[557.0, 275.0], [689.0, 467.0], [534.0, 629.0], [359.0, 475.0]] | 829.6105 | 2.6346 | 2.5056 | 8.2023 | 7.7851
Micro-Vickers¦711-081¦HV 1¦1¦row27 | Micro-Vickers¦SN=711-081¦scale=HV 1¦F=9.81N¦mag=20X | Micro-Vickers | 711-081_102_HMV_1_1_20X | 27 | 20X | 133.918 | 134.188 | 103.379 | 102.963 | C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\indentations\711-081_102_HMV_1_1_20X.bmp | rejected | No suitable indentation contour detected
Micro-Vickers¦706-437¦HV 1¦1¦row30 | Micro-Vickers¦SN=706-437¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers | 706-437_307_HMV_1_1_50X | 30 | 50X | 77.7295 | 77.5165 | 306.858 | 308.543 | C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\indentations\706-437_307_HMV_1_1_50X.bmp | rejected | all configured image checks passed; excessive agreement difference | Vickers contour-corner | 11869.5 | 1060.7178 | 124 | 669 | 178 | 193 | 1.0843 | 0.34551 | 0.13257 | 0.35411 | 4 | 0.18164 | 84.6332 | 97.2197 | 62.2415 | 18.5157 | 19.4063 | 18.961 | 1.0481 | [[125.0, 754.0], [251.0, 669.0], [300.0, 827.0], [181.0, 855.0]] | 5159.8484 | -58.7685 | -58.5555 | -121.5599 | -121.3868
Micro-Vickers¦706-437¦HV 1¦2¦row31 | Micro-Vickers¦SN=706-437¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers | 706-437_307_HMV_1_2_50X | 31 | 50X | 75.6685 | 75.219 | 323.417 | 327.684 | C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\indentations\706-437_307_HMV_1_2_50X.bmp | rejected | all configured image checks passed; excessive agreement difference | Vickers contour-corner | 20113 | 1914.2986 | 671 | 532 | 331 | 247 | 1.3401 | 0.24601 | 0.068971 | 0.46308 | 4 | 0.34797 | 60.7222 | 144.6176 | 74 | 28.5916 | 25.7943 | 27.1929 | 1.1084 | [[982.0, 532.0], [935.0, 716.0], [806.0, 766.0], [671.0, 707.0]] | 2508.6993 | -48.4756 | -48.0261 | -94.2541 | -93.79
Micro-Vickers¦711-246¦HV 1¦1¦row33 | Micro-Vickers¦SN=711-246¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers | 711-246_502_HMV_1_1_50X | 33 | 50X | 60.902 | 60.9225 | 499.858 | 499.513 | C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\indentations\711-246_502_HMV_1_1_50X.bmp | rejected | all configured image checks passed; excessive agreement difference | Vickers contour-corner | 401759.5 | 8218.9254 | 74 | 56 | 910 | 787 | 1.1563 | 0.56098 | 0.074739 | 0.9974 | 4 | 0.043619 | 50.4246 | 136.9164 | 29.7321 | 84.5449 | 89.6368 | 87.0909 | 1.0602 | [[713.0, 63.0], [971.0, 340.0], [325.0, 837.0], [74.0, 535.0]] | 244.5768 | 26.1889 | 26.1684 | 35.3921 | 35.3595
Micro-Vickers¦711-246¦HV 1¦3¦row35 | Micro-Vickers¦SN=711-246¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers | 711-246_502_HMV_1_3_50X | 35 | 50X | 60.0825 | 59.969 | 513.587 | 515.533 | C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\indentations\711-246_502_HMV_1_3_50X.bmp | rejected | all configured image checks passed; excessive agreement difference | Vickers contour-corner | 249872.5 | 5310.6227 | 122 | 47 | 757 | 689 | 1.0987 | 0.47907 | 0.11134 | 0.90925 | 4 | 0.11867 | 45.7685 | 185.8577 | 90.2441 | 56.0546 | 77.3293 | 66.6919 | 1.3795 | [[323.0, 192.0], [732.0, 47.0], [876.0, 346.0], [122.0, 552.0]] | 417.0748 | 6.6094 | 6.7229 | 10.4271 | 10.6156
Micro-Vickers¦710-106¦HV 1¦1¦row36 | Micro-Vickers¦SN=710-106¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers | 710-106_708_HMV_1_1_50X | 36 | 50X | 50.668 | 50.3055 | 722.175 | 732.635 | C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\indentations\710-106_708_HMV_1_1_50X.bmp | rejected | all configured image checks passed; excessive agreement difference | Vickers contour-corner | 173078 | 3637.7586 | 265 | 111 | 568 | 604 | 1.0634 | 0.50449 | 0.16436 | 0.94105 | 4 | 0.079331 | 50.0177 | 166.8095 | 50.9215 | 54.3409 | 60.3849 | 57.3629 | 1.1112 | [[295.0, 296.0], [586.0, 111.0], [832.0, 442.0], [768.0, 702.0]] | 563.7646 | 6.6949 | 7.0574 | 12.3945 | 13.1096
Micro-Vickers¦709-293¦HV 1¦1¦row39 | Micro-Vickers¦SN=709-293¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers | 709-293_904_HMV_1_1_50X | 39 | 50X | 45.101 | 44.291 | 911.46 | 945.102 | C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\indentations\709-293_904_HMV_1_1_50X.bmp | conditionally reliable | all configured image checks passed; elevated agreement difference | Vickers contour-corner | 113963.5 | 1625.124 | 302 | 206 | 465 | 466 | 1.0022 | 0.52593 | 0.54225 | 0.98397 | 4 | 0.055337 | 80.0114 | 109.1556 | 26.6833 | 45.2114 | 45.4101 | 45.3107 | 1.0044 | [[553.0, 208.0], [766.0, 436.0], [553.0, 671.0], [302.0, 467.0]] | 903.5632 | 0.20972 | 1.0197 | 0.46393 | 2.2761
Micro-Vickers¦711-081¦HV 2¦1¦row42 | Micro-Vickers¦SN=711-081¦scale=HV 2¦F=19.62N¦mag=20X | Micro-Vickers | 711-081_102_HMV_2_1_20X | 42 | 20X | 189.9425 | 190.1445 | 102.777 | 102.558 | C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\indentations\711-081_102_HMV_2_1_20X.bmp | rejected | No suitable indentation contour detected
Micro-Vickers¦706-437¦HV 2¦1¦row45 | Micro-Vickers¦SN=706-437¦scale=HV 2¦F=19.62N¦mag=20X | Micro-Vickers | 706-437_307_HMV_2_1_50X | 45 | 20X | 108.5715 | 108.619 | 314.563 | 314.288 | C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\indentations\706-437_307_HMV_2_1_50X.bmp | rejected | all configured image checks passed; excessive agreement difference | Vickers contour-corner | 1219.5 | 226.1838 | 978 | 146 | 66 | 43 | 1.5349 | 0.4297 | 0.29955 | 0.19338 | 4 | 0.17614 | 58.4674 | 129.3703 | 69.0652 | 14.6423 | 15.1003 | 14.8713 | 1.0313 | [[991.0, 152.0], [1031.0, 152.0], [1043.0, 182.0], [978.0, 184.0]] | 16776.194 | -93.7002 | -93.7477 | -151.8116 | -151.8301
Micro-Vickers¦711-246¦HV 2¦1¦row48 | Micro-Vickers¦SN=711-246¦scale=HV 2¦F=19.62N¦mag=50X | Micro-Vickers | 711-246_502_HMV_2_1_50X | 48 | 50X | 84.672 | 84.563 | 517.202 | 518.536 | C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\indentations\711-246_502_HMV_2_1_50X.bmp | rejected | opposite-corner consistency failed; unstable radial corner geometry; excessive agreement difference | Vickers contour-corner | 2256 | 345.8234 | 234 | 670 | 85 | 109 | 1.2824 | 0.2435 | 0.23705 | 0.46022 | 4 | 0.59466 | 55.8616 | 117.9859 | 78.5238 | 13.3604 | 3.333 | 8.3467 | 4.0086 | [[234.0, 670.0], [284.0, 717.0], [318.0, 778.0], [255.0, 735.0]] | 53255.2584 | -76.3253 | -76.2163 | -164.1075 | -164.0654
Micro-Vickers¦710-106¦HV 2¦1¦row51 | Micro-Vickers¦SN=710-106¦scale=HV 2¦F=19.62N¦mag=50X | Micro-Vickers | 710-106_708_HMV_2_1_50X | 51 | 50X | 70.8015 | 70.76 | 739.699 | 740.567 | C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\indentations\710-106_708_HMV_2_1_50X.bmp | rejected | all configured image checks passed; excessive agreement difference | Vickers contour-corner | 316965.5 | 5489.8633 | 150 | 61 | 779 | 731 | 1.0657 | 0.55662 | 0.13216 | 0.96479 | 4 | 0.078055 | 51.9888 | 172.1618 | 49.3356 | 76.2139 | 74.9012 | 75.5576 | 1.0175 | [[833.0, 122.0], [917.0, 586.0], [431.0, 791.0], [152.0, 530.0]] | 649.882 | 4.7561 | 4.7976 | 6.4992 | 6.5578
Micro-Vickers¦709-293¦HV 2¦1¦row54 | Micro-Vickers¦SN=709-293¦scale=HV 2¦F=19.62N¦mag=50X | Micro-Vickers | 709-293_904_HMV_2_1_50X | 54 | 50X | 63.6595 | 63.747 | 914.984 | 912.473 | C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\indentations\709-293_904_HMV_2_1_50X.bmp | conditionally reliable | all configured image checks passed; elevated agreement difference | Vickers contour-corner | 225888 | 2597.4987 | 191 | 107 | 675 | 670 | 1.0075 | 0.49948 | 0.42072 | 0.99131 | 4 | 0.11441 | 72.6869 | 106.4218 | 30.5941 | 65.5434 | 65.1344 | 65.3389 | 1.0063 | [[191.0, 441.0], [480.0, 110.0], [862.0, 424.0], [517.0, 776.0]] | 869.0549 | 1.6794 | 1.5919 | 2.6037 | 2.4664
Brinell¦293-250¦HBW 1/5¦1¦row3 | Brinell¦SN=293-250¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | Brinell | 100_HBW_1_5_1_20X | 3 | 20X | 210.949 | 212.334 | 141.497 | 139.636 | C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\indentations\100_HBW_1_5_1_20X.bmp | rejected | incomplete contour touches image boundary; large radial residual variation; excessive agreement difference | Brinell contour-diameter | 944225 | 3900 | 0 | 0 | 1056 | 896 | 1.1786 | 0.99794 | 0.78011 | 1 | [5 empty] | 257.561 | 218.5366 | 238.0488 |  |  | 110.7975 | 27.0998 | 25.7148 | 12.0712 | 11.4191 | 168.7186 | 16.8198 | 0.12344 | 307.7612 | 255.7581 | 1.2033 | 118.0142 | 113.6968 | 257.3171 | 336.3069 | 218.2927 | 336.3069 | 257.3171 | 331.9895 | 317.7055 | 218.2927 | 317.7055 | 331.9895
Brinell¦294-367¦HBW 1/5¦1¦row6 | Brinell¦SN=294-367¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | Brinell | 300_HBW_1_5_1_20X | 6 | 20X | 154.623 | 155.5175 | 264.758 | 261.701 | C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\indentations\300_HBW_1_5_1_20X.bmp | rejected | incomplete contour touches image boundary; large radial residual variation; excessive agreement difference | Brinell contour-diameter | 944225 | 3900 | 0 | 0 | 1056 | 896 | 1.1786 | 0.99794 | 0.78011 | 1 | [5 empty] | 257.561 | 218.5366 | 238.0488 |  |  | 110.7975 | 83.4258 | 82.5313 | 42.4914 | 41.9402 | 168.7186 | 16.8198 | 0.12344 | 307.7612 | 255.7581 | 1.2033 | 118.0142 | 113.6968 | 257.3171 | 336.3069 | 218.2927 | 336.3069 | 257.3171 | 331.9895 | 317.7055 | 218.2927 | 317.7055 | 331.9895
Brinell¦290-736¦HBW 1/5¦1¦row9 | Brinell¦SN=290-736¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | Brinell | 600_HBW_1_5_1_20X | 9 | 20X | 115.937 | 117.397 | 472.178 | 460.467 | C:\Users\Korisnik\Desktop\RMS-TracInd-analysis\images\indentations\600_HBW_1_5_1_20X.bmp | rejected | incomplete contour touches image boundary; large radial residual variation; excessive agreement difference | Brinell contour-diameter | 944225 | 3900 | 0 | 0 | 1056 | 896 | 1.1786 | 0.99794 | 0.78011 | 1 | [5 empty] | 257.561 | 218.5366 | 238.0488 |  |  | 110.7975 | 122.1118 | 120.6518 | 68.9925 | 67.8876 | 168.7186 | 16.8198 | 0.12344 | 307.7612 | 255.7581 | 1.2033 | 118.0142 | 113.6968 | 257.3171 | 336.3069 | 218.2927 | 336.3069 | 257.3171 | 331.9895 | 317.7055 | 218.2927 | 317.7055 | 331.9895

## Image_Quality
record_id | series_id | method | picture_id | magnification | image_status | classification_reason | edge_strength | diagonal_ratio | corner_radius_cv | min_corner_angle_deg | circularity | ellipse_axis_ratio | radial_cv | diameter_direction_range_4_um | python_minus_manual_relative_percent | python_minus_automatic_relative_percent
Vickers¦293-250¦HV 3¦1¦row3 | Vickers¦SN=293-250¦scale=HV 3¦F=29.43N¦mag=20X | Vickers | 293-250_160_HV_3_1_20X | 20X | conditionally reliable | all configured image checks passed; elevated agreement difference | 132.8345 | 1.0143 | 0.086675 | 48.0967 | 0.12553 |  |  |  | 3.8252 | 3.5885
Vickers¦294-367¦HV 3¦1¦row6 | Vickers¦SN=294-367¦scale=HV 3¦F=29.43N¦mag=20X | Vickers | 294-367_300_HV_3_1_20X | 20X | rejected | all configured image checks passed; excessive agreement difference | 61.6117 | 1.1733 | 0.058921 | 50.1744 | 0.13784 |  |  |  | 28.2575 | 28.1924
Vickers¦294-481¦HV 3¦1¦row9 | Vickers¦SN=294-481¦scale=HV 3¦F=29.43N¦mag=20X | Vickers | 294_481_503_HV_3_1_20X | 20X | rejected | all configured image checks passed; excessive agreement difference | 22.6716 | 1.0307 | 0.066701 | 67.0558 | 0.12992 |  |  |  | 47.1945 | 46.855
Vickers¦216-610¦HV 3¦1¦row12 | Vickers¦SN=216-610¦scale=HV 3¦F=29.43N¦mag=50X | Vickers | 216-610_701_HV_3_1_50X | 50X | rejected | all configured image checks passed; excessive agreement difference | 90.3264 | 1.0048 | 0.13693 | 51.9658 | 0.27539 |  |  |  | -167.7697 | -167.5899
Vickers¦221-101¦HV 3¦1¦row15 | Vickers¦SN=221-101¦scale=HV 3¦F=29.43N¦mag=50X | Vickers | 221-101_898_HV_3_1_50X | 50X | reliable | all configured image checks passed | 26.0768 | 1.0006 | 0.055869 | 80.9002 | 0.44285 |  |  |  | 1.6616 | 1.9643
Vickers¦294-481¦HV 20¦1¦row24 | Vickers¦SN=294-481¦scale=HV 20¦F=196.2N¦mag=20X | Vickers | 294_481_503_HV_20_1_20X | 20X | rejected | No suitable indentation contour detected
Vickers¦216-610¦HV 20¦1¦row27 | Vickers¦SN=216-610¦scale=HV 20¦F=196.2N¦mag=20X | Vickers | 216-610_701_HV_20_1_20X | 20X | reliable | all configured image checks passed | 171.7323 | 1.0034 | 0.057734 | 82.0894 | 0.58571 |  |  |  | 0.058495 | 0.059593
Vickers¦221-101¦HV 20¦1¦row30 | Vickers¦SN=221-101¦scale=HV 20¦F=196.2N¦mag=20X | Vickers | 221-101_898_HV_20_1_20X | 20X | conditionally reliable | all configured image checks passed; elevated agreement difference | 17.0294 | 1.0033 | 0.044083 | 82.2368 | 0.49639 |  |  |  | 2.3533 | 2.3168
Micro-Vickers¦710-404¦HV 0.05¦1¦row3 | Micro-Vickers¦SN=710-404¦scale=HV 0.05¦F=0.4905N¦mag=50X | Micro-Vickers | 710-404_39.6_HMV_0.05_1_50X | 50X | rejected | all configured image checks passed; excessive agreement difference | 96.7884 | 1.2442 | 0.27271 | 75.7637 | 0.27822 |  |  |  | 18.864 | 18.4993
Micro-Vickers¦711-081¦HV 0.05¦1¦row6 | Micro-Vickers¦SN=711-081¦scale=HV 0.05¦F=0.4905N¦mag=50X | Micro-Vickers | 711-081_102_HMV_0.05_1_50X | 50X | rejected | all configured image checks passed; excessive agreement difference | 55.118 | 1.0577 | 0.20764 | 58.1812 | 0.069679 |  |  |  | 55.1017 | 54.9983
Micro-Vickers¦710-655¦HV 0.05¦1¦row9 | Micro-Vickers¦SN=710-655¦scale=HV 0.05¦F=0.4905N¦mag=50X | Micro-Vickers | 710-655_197_HMV_0.05_1_50X | 50X | rejected | all configured image checks passed; excessive agreement difference | 100.0649 | 1.0615 | 0.18613 | 47.5464 | 0.22744 |  |  |  | 26.158 | 25.866
Micro-Vickers¦711-081¦HV 0.5¦1¦row12 | Micro-Vickers¦SN=711-081¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers | 711-081_102_HMV_0.5_1_50X | 50X | rejected | No suitable indentation contour detected
Micro-Vickers¦706-437¦HV 0.5¦1¦row15 | Micro-Vickers¦SN=706-437¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers | 706-437_307_HMV_0.5_1_50X | 50X | conditionally reliable | all configured image checks passed; elevated agreement difference | 105.8017 | 1.0631 | 0.062965 | 54.6586 | 0.23644 |  |  |  | -2.454 | -2.2419
Micro-Vickers¦711-246¦HV 0.5¦1¦row18 | Micro-Vickers¦SN=711-246¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers | 711-246_502_HMV_0.5_1_50X | 50X | rejected | all configured image checks passed; excessive agreement difference | 29.2913 | 1.1028 | 0.15024 | 67.9303 | 0.054972 |  |  |  | 45.1883 | 44.8304
Micro-Vickers¦710-106¦HV 0.5¦1¦row21 | Micro-Vickers¦SN=710-106¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers | 710-106_708_HMV_0.5_1_50X | 50X | rejected | all configured image checks passed; excessive agreement difference | 30.9515 | 1.1269 | 0.046667 | 45.0231 | 0.38084 |  |  |  | -10.3807 | -9.5944
Micro-Vickers¦709-293¦HV 0.5¦1¦row24 | Micro-Vickers¦SN=709-293¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers | 709-293_904_HMV_0.5_1_50X | 50X | rejected | all configured image checks passed; excessive agreement difference | 35.9303 | 1.0747 | 0.064842 | 77.1666 | 0.58406 |  |  |  | 8.2023 | 7.7851
Micro-Vickers¦711-081¦HV 1¦1¦row27 | Micro-Vickers¦SN=711-081¦scale=HV 1¦F=9.81N¦mag=20X | Micro-Vickers | 711-081_102_HMV_1_1_20X | 20X | rejected | No suitable indentation contour detected
Micro-Vickers¦706-437¦HV 1¦1¦row30 | Micro-Vickers¦SN=706-437¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers | 706-437_307_HMV_1_1_50X | 50X | rejected | all configured image checks passed; excessive agreement difference | 62.2415 | 1.0481 | 0.18164 | 84.6332 | 0.13257 |  |  |  | -121.5599 | -121.3868
Micro-Vickers¦706-437¦HV 1¦2¦row31 | Micro-Vickers¦SN=706-437¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers | 706-437_307_HMV_1_2_50X | 50X | rejected | all configured image checks passed; excessive agreement difference | 74 | 1.1084 | 0.34797 | 60.7222 | 0.068971 |  |  |  | -94.2541 | -93.79
Micro-Vickers¦711-246¦HV 1¦1¦row33 | Micro-Vickers¦SN=711-246¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers | 711-246_502_HMV_1_1_50X | 50X | rejected | all configured image checks passed; excessive agreement difference | 29.7321 | 1.0602 | 0.043619 | 50.4246 | 0.074739 |  |  |  | 35.3921 | 35.3595
Micro-Vickers¦711-246¦HV 1¦3¦row35 | Micro-Vickers¦SN=711-246¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers | 711-246_502_HMV_1_3_50X | 50X | rejected | all configured image checks passed; excessive agreement difference | 90.2441 | 1.3795 | 0.11867 | 45.7685 | 0.11134 |  |  |  | 10.4271 | 10.6156
Micro-Vickers¦710-106¦HV 1¦1¦row36 | Micro-Vickers¦SN=710-106¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers | 710-106_708_HMV_1_1_50X | 50X | rejected | all configured image checks passed; excessive agreement difference | 50.9215 | 1.1112 | 0.079331 | 50.0177 | 0.16436 |  |  |  | 12.3945 | 13.1096
Micro-Vickers¦709-293¦HV 1¦1¦row39 | Micro-Vickers¦SN=709-293¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers | 709-293_904_HMV_1_1_50X | 50X | conditionally reliable | all configured image checks passed; elevated agreement difference | 26.6833 | 1.0044 | 0.055337 | 80.0114 | 0.54225 |  |  |  | 0.46393 | 2.2761
Micro-Vickers¦711-081¦HV 2¦1¦row42 | Micro-Vickers¦SN=711-081¦scale=HV 2¦F=19.62N¦mag=20X | Micro-Vickers | 711-081_102_HMV_2_1_20X | 20X | rejected | No suitable indentation contour detected
Micro-Vickers¦706-437¦HV 2¦1¦row45 | Micro-Vickers¦SN=706-437¦scale=HV 2¦F=19.62N¦mag=20X | Micro-Vickers | 706-437_307_HMV_2_1_50X | 20X | rejected | all configured image checks passed; excessive agreement difference | 69.0652 | 1.0313 | 0.17614 | 58.4674 | 0.29955 |  |  |  | -151.8116 | -151.8301
Micro-Vickers¦711-246¦HV 2¦1¦row48 | Micro-Vickers¦SN=711-246¦scale=HV 2¦F=19.62N¦mag=50X | Micro-Vickers | 711-246_502_HMV_2_1_50X | 50X | rejected | opposite-corner consistency failed; unstable radial corner geometry; excessive agreement difference | 78.5238 | 4.0086 | 0.59466 | 55.8616 | 0.23705 |  |  |  | -164.1075 | -164.0654
Micro-Vickers¦710-106¦HV 2¦1¦row51 | Micro-Vickers¦SN=710-106¦scale=HV 2¦F=19.62N¦mag=50X | Micro-Vickers | 710-106_708_HMV_2_1_50X | 50X | rejected | all configured image checks passed; excessive agreement difference | 49.3356 | 1.0175 | 0.078055 | 51.9888 | 0.13216 |  |  |  | 6.4992 | 6.5578
Micro-Vickers¦709-293¦HV 2¦1¦row54 | Micro-Vickers¦SN=709-293¦scale=HV 2¦F=19.62N¦mag=50X | Micro-Vickers | 709-293_904_HMV_2_1_50X | 50X | conditionally reliable | all configured image checks passed; elevated agreement difference | 30.5941 | 1.0063 | 0.11441 | 72.6869 | 0.42072 |  |  |  | 2.6037 | 2.4664
Brinell¦293-250¦HBW 1/5¦1¦row3 | Brinell¦SN=293-250¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | Brinell | 100_HBW_1_5_1_20X | 20X | rejected | incomplete contour touches image boundary; large radial residual variation; excessive agreement difference |  |  |  |  | 0.78011 | 1.2033 | 0.12344 | 118.0142 | 12.0712 | 11.4191
Brinell¦294-367¦HBW 1/5¦1¦row6 | Brinell¦SN=294-367¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | Brinell | 300_HBW_1_5_1_20X | 20X | rejected | incomplete contour touches image boundary; large radial residual variation; excessive agreement difference |  |  |  |  | 0.78011 | 1.2033 | 0.12344 | 118.0142 | 42.4914 | 41.9402
Brinell¦290-736¦HBW 1/5¦1¦row9 | Brinell¦SN=290-736¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | Brinell | 600_HBW_1_5_1_20X | 20X | rejected | incomplete contour touches image boundary; large radial residual variation; excessive agreement difference |  |  |  |  | 0.78011 | 1.2033 | 0.12344 | 118.0142 | 68.9925 | 67.8876

## Classification
record_id | series_id | method | picture_id | magnification | suitability_class | reason
Vickers¦293-250¦HV 3¦1¦row3 | Vickers¦SN=293-250¦scale=HV 3¦F=29.43N¦mag=20X | Vickers | 293-250_160_HV_3_1_20X | 20X | conditionally suitable/requires review | all configured image checks passed; elevated agreement difference
Vickers¦294-367¦HV 3¦1¦row6 | Vickers¦SN=294-367¦scale=HV 3¦F=29.43N¦mag=20X | Vickers | 294-367_300_HV_3_1_20X | 20X | unsuitable for unattended automatic measurement | all configured image checks passed; excessive agreement difference; excessive Python-automatic difference
Vickers¦294-481¦HV 3¦1¦row9 | Vickers¦SN=294-481¦scale=HV 3¦F=29.43N¦mag=20X | Vickers | 294_481_503_HV_3_1_20X | 20X | unsuitable for unattended automatic measurement | all configured image checks passed; excessive agreement difference; excessive Python-automatic difference
Vickers¦216-610¦HV 3¦1¦row12 | Vickers¦SN=216-610¦scale=HV 3¦F=29.43N¦mag=50X | Vickers | 216-610_701_HV_3_1_50X | 50X | unsuitable for unattended automatic measurement | all configured image checks passed; excessive agreement difference; excessive Python-automatic difference
Vickers¦221-101¦HV 3¦1¦row15 | Vickers¦SN=221-101¦scale=HV 3¦F=29.43N¦mag=50X | Vickers | 221-101_898_HV_3_1_50X | 50X | suitable for automatic measurement | configured image and agreement checks passed
Vickers¦294-481¦HV 20¦1¦row24 | Vickers¦SN=294-481¦scale=HV 20¦F=196.2N¦mag=20X | Vickers | 294_481_503_HV_20_1_20X | 20X | unsuitable for unattended automatic measurement | No suitable indentation contour detected
Vickers¦216-610¦HV 20¦1¦row27 | Vickers¦SN=216-610¦scale=HV 20¦F=196.2N¦mag=20X | Vickers | 216-610_701_HV_20_1_20X | 20X | suitable for automatic measurement | configured image and agreement checks passed
Vickers¦221-101¦HV 20¦1¦row30 | Vickers¦SN=221-101¦scale=HV 20¦F=196.2N¦mag=20X | Vickers | 221-101_898_HV_20_1_20X | 20X | conditionally suitable/requires review | all configured image checks passed; elevated agreement difference
Micro-Vickers¦710-404¦HV 0.05¦1¦row3 | Micro-Vickers¦SN=710-404¦scale=HV 0.05¦F=0.4905N¦mag=50X | Micro-Vickers | 710-404_39.6_HMV_0.05_1_50X | 50X | unsuitable for unattended automatic measurement | all configured image checks passed; excessive agreement difference; excessive Python-automatic difference
Micro-Vickers¦711-081¦HV 0.05¦1¦row6 | Micro-Vickers¦SN=711-081¦scale=HV 0.05¦F=0.4905N¦mag=50X | Micro-Vickers | 711-081_102_HMV_0.05_1_50X | 50X | unsuitable for unattended automatic measurement | all configured image checks passed; excessive agreement difference; excessive Python-automatic difference
Micro-Vickers¦710-655¦HV 0.05¦1¦row9 | Micro-Vickers¦SN=710-655¦scale=HV 0.05¦F=0.4905N¦mag=50X | Micro-Vickers | 710-655_197_HMV_0.05_1_50X | 50X | unsuitable for unattended automatic measurement | all configured image checks passed; excessive agreement difference; excessive Python-automatic difference
Micro-Vickers¦711-081¦HV 0.5¦1¦row12 | Micro-Vickers¦SN=711-081¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers | 711-081_102_HMV_0.5_1_50X | 50X | unsuitable for unattended automatic measurement | No suitable indentation contour detected
Micro-Vickers¦706-437¦HV 0.5¦1¦row15 | Micro-Vickers¦SN=706-437¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers | 706-437_307_HMV_0.5_1_50X | 50X | conditionally suitable/requires review | all configured image checks passed; elevated agreement difference
Micro-Vickers¦711-246¦HV 0.5¦1¦row18 | Micro-Vickers¦SN=711-246¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers | 711-246_502_HMV_0.5_1_50X | 50X | unsuitable for unattended automatic measurement | all configured image checks passed; excessive agreement difference; excessive Python-automatic difference
Micro-Vickers¦710-106¦HV 0.5¦1¦row21 | Micro-Vickers¦SN=710-106¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers | 710-106_708_HMV_0.5_1_50X | 50X | unsuitable for unattended automatic measurement | all configured image checks passed; excessive agreement difference; excessive Python-automatic difference
Micro-Vickers¦709-293¦HV 0.5¦1¦row24 | Micro-Vickers¦SN=709-293¦scale=HV 0.5¦F=4.905N¦mag=50X | Micro-Vickers | 709-293_904_HMV_0.5_1_50X | 50X | unsuitable for unattended automatic measurement | all configured image checks passed; excessive agreement difference; excessive Python-automatic difference
Micro-Vickers¦711-081¦HV 1¦1¦row27 | Micro-Vickers¦SN=711-081¦scale=HV 1¦F=9.81N¦mag=20X | Micro-Vickers | 711-081_102_HMV_1_1_20X | 20X | unsuitable for unattended automatic measurement | No suitable indentation contour detected
Micro-Vickers¦706-437¦HV 1¦1¦row30 | Micro-Vickers¦SN=706-437¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers | 706-437_307_HMV_1_1_50X | 50X | unsuitable for unattended automatic measurement | all configured image checks passed; excessive agreement difference; excessive Python-automatic difference
Micro-Vickers¦706-437¦HV 1¦2¦row31 | Micro-Vickers¦SN=706-437¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers | 706-437_307_HMV_1_2_50X | 50X | unsuitable for unattended automatic measurement | all configured image checks passed; excessive agreement difference; excessive Python-automatic difference
Micro-Vickers¦711-246¦HV 1¦1¦row33 | Micro-Vickers¦SN=711-246¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers | 711-246_502_HMV_1_1_50X | 50X | unsuitable for unattended automatic measurement | all configured image checks passed; excessive agreement difference; excessive Python-automatic difference
Micro-Vickers¦711-246¦HV 1¦3¦row35 | Micro-Vickers¦SN=711-246¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers | 711-246_502_HMV_1_3_50X | 50X | unsuitable for unattended automatic measurement | all configured image checks passed; excessive agreement difference; excessive Python-automatic difference
Micro-Vickers¦710-106¦HV 1¦1¦row36 | Micro-Vickers¦SN=710-106¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers | 710-106_708_HMV_1_1_50X | 50X | unsuitable for unattended automatic measurement | all configured image checks passed; excessive agreement difference; excessive Python-automatic difference
Micro-Vickers¦709-293¦HV 1¦1¦row39 | Micro-Vickers¦SN=709-293¦scale=HV 1¦F=9.81N¦mag=50X | Micro-Vickers | 709-293_904_HMV_1_1_50X | 50X | conditionally suitable/requires review | all configured image checks passed; elevated agreement difference; ¦E_n¦ > 1 for corresponding series
Micro-Vickers¦711-081¦HV 2¦1¦row42 | Micro-Vickers¦SN=711-081¦scale=HV 2¦F=19.62N¦mag=20X | Micro-Vickers | 711-081_102_HMV_2_1_20X | 20X | unsuitable for unattended automatic measurement | No suitable indentation contour detected
Micro-Vickers¦706-437¦HV 2¦1¦row45 | Micro-Vickers¦SN=706-437¦scale=HV 2¦F=19.62N¦mag=20X | Micro-Vickers | 706-437_307_HMV_2_1_50X | 20X | unsuitable for unattended automatic measurement | all configured image checks passed; excessive agreement difference; excessive Python-automatic difference
Micro-Vickers¦711-246¦HV 2¦1¦row48 | Micro-Vickers¦SN=711-246¦scale=HV 2¦F=19.62N¦mag=50X | Micro-Vickers | 711-246_502_HMV_2_1_50X | 50X | unsuitable for unattended automatic measurement | opposite-corner consistency failed; unstable radial corner geometry; excessive agreement difference; excessive Python-automatic difference
Micro-Vickers¦710-106¦HV 2¦1¦row51 | Micro-Vickers¦SN=710-106¦scale=HV 2¦F=19.62N¦mag=50X | Micro-Vickers | 710-106_708_HMV_2_1_50X | 50X | unsuitable for unattended automatic measurement | all configured image checks passed; excessive agreement difference; excessive Python-automatic difference
Micro-Vickers¦709-293¦HV 2¦1¦row54 | Micro-Vickers¦SN=709-293¦scale=HV 2¦F=19.62N¦mag=50X | Micro-Vickers | 709-293_904_HMV_2_1_50X | 50X | conditionally suitable/requires review | all configured image checks passed; elevated agreement difference
Brinell¦293-250¦HBW 1/5¦1¦row3 | Brinell¦SN=293-250¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | Brinell | 100_HBW_1_5_1_20X | 20X | unsuitable for unattended automatic measurement | incomplete contour touches image boundary; large radial residual variation; excessive agreement difference; excessive Python-automatic difference
Brinell¦294-367¦HBW 1/5¦1¦row6 | Brinell¦SN=294-367¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | Brinell | 300_HBW_1_5_1_20X | 20X | unsuitable for unattended automatic measurement | incomplete contour touches image boundary; large radial residual variation; excessive agreement difference; excessive Python-automatic difference
Brinell¦290-736¦HBW 1/5¦1¦row9 | Brinell¦SN=290-736¦scale=HBW 1/5¦F=49.05N¦ball=1mm¦mag=20X | Brinell | 600_HBW_1_5_1_20X | 20X | unsuitable for unattended automatic measurement | incomplete contour touches image boundary; large radial residual variation; excessive agreement difference; excessive Python-automatic difference

## Exclusions
method | scope | identifier | reason
Image file | image | 100_HBW_10_3000_1_10X.bmp | Image file has no exact workbook Picture ID match
Image file | image | 100_HBW_10_500_1_10X.bmp | Image file has no exact workbook Picture ID match
Image file | image | 100_HBW_10_500_1_20X.bmp | Image file has no exact workbook Picture ID match
Image file | image | 100_HBW_1_30_1_10X.bmp | Image file has no exact workbook Picture ID match
Image file | image | 100_HK_0.05_1_50X.bmp | Image file has no exact workbook Picture ID match
Image file | image | 100_HK_1_1_10X.bmp | Image file has no exact workbook Picture ID match
Image file | image | 100_HK_2_1_10X.bmp | Image file has no exact workbook Picture ID match
Image file | image | 216-610_701_HV_100_1_10X.bmp | Image file has no exact workbook Picture ID match
Image file | image | 216-610_701_HV_50_1_10X.bmp | Image file has no exact workbook Picture ID match
Image file | image | 221-101_898_HV_100_1_10X.bmp | Image file has no exact workbook Picture ID match
Image file | image | 221-101_898_HV_50_1_10X.bmp | Image file has no exact workbook Picture ID match
Image file | image | 293-250_160_HV_20_1_10X.bmp | Image file has no exact workbook Picture ID match
Image file | image | 294-367_300_HV_20_1_10X.bmp | Image file has no exact workbook Picture ID match
Image file | image | 294-367_300_HV_50_1_10X.bmp | Image file has no exact workbook Picture ID match
Image file | image | 294_481_503_HV_50_1_10X.bmp | Image file has no exact workbook Picture ID match
Image file | image | 300_HBW_10_3000_1_10X.bmp | Image file has no exact workbook Picture ID match
Image file | image | 300_HBW_10_500_1_10X.bmp | Image file has no exact workbook Picture ID match
Image file | image | 300_HBW_1_30_1_10X.bmp | Image file has no exact workbook Picture ID match
Image file | image | 300_HBW_1_30_1_20X.bmp | Image file has no exact workbook Picture ID match
Image file | image | 500_HK_0.05_1_50X.bmp | Image file has no exact workbook Picture ID match
Image file | image | 500_HK_1_1_20X.bmp | Image file has no exact workbook Picture ID match
Image file | image | 500_HK_2_1_20X.bmp | Image file has no exact workbook Picture ID match
Image file | image | 500_KH_1_1_20X.bmp | Image file has no exact workbook Picture ID match
Image file | image | 600_HBW_10_3000_1_10X.bmp | Image file has no exact workbook Picture ID match
Image file | image | 600_HBW_10_3000_1_20X.bmp | Image file has no exact workbook Picture ID match
Image file | image | 600_HBW_10_500_1_10X.bmp | Image file has no exact workbook Picture ID match
Image file | image | 600_HBW_1_30_1_10X.bmp | Image file has no exact workbook Picture ID match
Image file | image | 600_HBW_1_30_1_20X.bmp | Image file has no exact workbook Picture ID match
Image file | image | 900_HK_0.05_1_50X.bmp | Image file has no exact workbook Picture ID match
Image file | image | 900_HK_2_1_20X.bmp | Image file has no exact workbook Picture ID match

## Source_Files
folder | file | suffix | size_bytes
data | Hardness_measurement_results.xlsx | .xlsx | 54641
docs | 22RPT01 Tracind BVK-H Annex 1 v1.0.pdf | .pdf | 1855442
docs | APPLICATION.docx | .docx | 17882
docs | Research Mobility Support (RMS) Applicants Guide.pdf | .pdf | 399497
docs | RMS Sampling.docx | .docx | 20319
images\indentations | 100_HBW_10_3000_1_10X.bmp | .bmp | 523535
images\indentations | 100_HBW_10_500_1_10X.bmp | .bmp | 669740
images\indentations | 100_HBW_10_500_1_20X.bmp | .bmp | 744589
images\indentations | 100_HBW_1_30_1_10X.bmp | .bmp | 423503
images\indentations | 100_HBW_1_5_1_20X.bmp | .bmp | 650148
images\indentations | 100_HK_0.05_1_50X.bmp | .bmp | 580039
images\indentations | 100_HK_1_1_10X.bmp | .bmp | 604413
images\indentations | 100_HK_2_1_10X.bmp | .bmp | 609152
images\indentations | 216-610_701_HV_100_1_10X.bmp | .bmp | 416864
images\indentations | 216-610_701_HV_20_1_20X.bmp | .bmp | 494540
images\indentations | 216-610_701_HV_3_1_50X.bmp | .bmp | 590707
images\indentations | 216-610_701_HV_50_1_10X.bmp | .bmp | 498921
images\indentations | 221-101_898_HV_100_1_10X.bmp | .bmp | 414809
images\indentations | 221-101_898_HV_20_1_20X.bmp | .bmp | 457953
images\indentations | 221-101_898_HV_3_1_50X.bmp | .bmp | 535056
images\indentations | 221-101_898_HV_50_1_10X.bmp | .bmp | 465090
images\indentations | 293-250_160_HV_20_1_10X.bmp | .bmp | 490048
images\indentations | 293-250_160_HV_3_1_20X.bmp | .bmp | 571446
images\indentations | 294-367_300_HV_20_1_10X.bmp | .bmp | 560350
images\indentations | 294-367_300_HV_3_1_20X.bmp | .bmp | 524415
images\indentations | 294-367_300_HV_50_1_10X.bmp | .bmp | 399945
images\indentations | 294_481_503_HV_20_1_20X.bmp | .bmp | 413022
images\indentations | 294_481_503_HV_3_1_20X.bmp | .bmp | 485254
images\indentations | 294_481_503_HV_50_1_10X.bmp | .bmp | 526979
images\indentations | 300_HBW_10_3000_1_10X.bmp | .bmp | 611499
images\indentations | 300_HBW_10_500_1_10X.bmp | .bmp | 870625
images\indentations | 300_HBW_1_30_1_10X.bmp | .bmp | 553904
images\indentations | 300_HBW_1_30_1_20X.bmp | .bmp | 584332
images\indentations | 300_HBW_1_5_1_20X.bmp | .bmp | 569647
images\indentations | 500_HK_0.05_1_50X.bmp | .bmp | 486445
images\indentations | 500_HK_1_1_20X.bmp | .bmp | 520824
images\indentations | 500_HK_2_1_20X.bmp | .bmp | 528617
images\indentations | 500_KH_1_1_20X.bmp | .bmp | 520824
images\indentations | 600_HBW_10_3000_1_10X.bmp | .bmp | 737537
images\indentations | 600_HBW_10_3000_1_20X.bmp | .bmp | 762695
images\indentations | 600_HBW_10_500_1_10X.bmp | .bmp | 763574
images\indentations | 600_HBW_1_30_1_10X.bmp | .bmp | 626145
images\indentations | 600_HBW_1_30_1_20X.bmp | .bmp | 604648
images\indentations | 600_HBW_1_5_1_20X.bmp | .bmp | 524996
images\indentations | 706-437_307_HMV_0.5_1_50X.bmp | .bmp | 572225
images\indentations | 706-437_307_HMV_1_1_50X.bmp | .bmp | 638821
images\indentations | 706-437_307_HMV_1_2_50X.bmp | .bmp | 633955
images\indentations | 706-437_307_HMV_2_1_50X.bmp | .bmp | 641314
images\indentations | 709-293_904_HMV_0.5_1_50X.bmp | .bmp | 478406
images\indentations | 709-293_904_HMV_1_1_50X.bmp | .bmp | 486818
images\indentations | 709-293_904_HMV_2_1_50X.bmp | .bmp | 531106
images\indentations | 710-106_708_HMV_0.5_1_50X.bmp | .bmp | 497407
images\indentations | 710-106_708_HMV_1_1_50X.bmp | .bmp | 553336
images\indentations | 710-106_708_HMV_2_1_50X.bmp | .bmp | 584990
images\indentations | 710-404_39.6_HMV_0.05_1_50X.bmp | .bmp | 573923
images\indentations | 710-655_197_HMV_0.05_1_50X.bmp | .bmp | 602669
images\indentations | 711-081_102_HMV_0.05_1_50X.bmp | .bmp | 563576
images\indentations | 711-081_102_HMV_0.5_1_50X.bmp | .bmp | 631314
images\indentations | 711-081_102_HMV_1_1_20X.bmp | .bmp | 636930
images\indentations | 711-081_102_HMV_2_1_20X.bmp | .bmp | 637374
images\indentations | 711-246_502_HMV_0.5_1_50X.bmp | .bmp | 532149
images\indentations | 711-246_502_HMV_1_1_50X.bmp | .bmp | 562900
images\indentations | 711-246_502_HMV_1_3_50X.bmp | .bmp | 590496
images\indentations | 711-246_502_HMV_2_1_50X.bmp | .bmp | 633983
images\indentations | 900_HK_0.05_1_50X.bmp | .bmp | 456048
images\indentations | 900_HK_2_1_20X.bmp | .bmp | 492591
images\calibration | stage_micrometer_0.002mm_20x.bmp | .bmp | 495991
images\calibration | stage_micrometer_0.002mm_50x.bmp | .bmp | 418798

## Image_Dimensions
image_kind | file | width_px | height_px | mode
calibration | stage_micrometer_0.002mm_20x.bmp | 1056 | 896 | P
calibration | stage_micrometer_0.002mm_50x.bmp | 1056 | 896 | P
indentation | 100_HBW_10_3000_1_10X.bmp | 1056 | 896 | P
indentation | 100_HBW_10_500_1_10X.bmp | 1056 | 896 | P
indentation | 100_HBW_10_500_1_20X.bmp | 1056 | 896 | P
indentation | 100_HBW_1_30_1_10X.bmp | 1056 | 896 | P
indentation | 100_HBW_1_5_1_20X.bmp | 1056 | 896 | P
indentation | 100_HK_0.05_1_50X.bmp | 1056 | 896 | P
indentation | 100_HK_1_1_10X.bmp | 1056 | 896 | P
indentation | 100_HK_2_1_10X.bmp | 1056 | 896 | P
indentation | 216-610_701_HV_100_1_10X.bmp | 1056 | 896 | P
indentation | 216-610_701_HV_20_1_20X.bmp | 1056 | 896 | P
indentation | 216-610_701_HV_3_1_50X.bmp | 1056 | 896 | P
indentation | 216-610_701_HV_50_1_10X.bmp | 1056 | 896 | P
indentation | 221-101_898_HV_100_1_10X.bmp | 1056 | 896 | P
indentation | 221-101_898_HV_20_1_20X.bmp | 1056 | 896 | P
indentation | 221-101_898_HV_3_1_50X.bmp | 1056 | 896 | P
indentation | 221-101_898_HV_50_1_10X.bmp | 1056 | 896 | P
indentation | 293-250_160_HV_20_1_10X.bmp | 1056 | 896 | P
indentation | 293-250_160_HV_3_1_20X.bmp | 1056 | 896 | P
indentation | 294-367_300_HV_20_1_10X.bmp | 1056 | 896 | P
indentation | 294-367_300_HV_3_1_20X.bmp | 1056 | 896 | P
indentation | 294-367_300_HV_50_1_10X.bmp | 1056 | 896 | P
indentation | 294_481_503_HV_20_1_20X.bmp | 1056 | 896 | P
indentation | 294_481_503_HV_3_1_20X.bmp | 1056 | 896 | P
indentation | 294_481_503_HV_50_1_10X.bmp | 1056 | 896 | P
indentation | 300_HBW_10_3000_1_10X.bmp | 1056 | 896 | P
indentation | 300_HBW_10_500_1_10X.bmp | 1056 | 896 | P
indentation | 300_HBW_1_30_1_10X.bmp | 1056 | 896 | P
indentation | 300_HBW_1_30_1_20X.bmp | 1056 | 896 | P
indentation | 300_HBW_1_5_1_20X.bmp | 1056 | 896 | P
indentation | 500_HK_0.05_1_50X.bmp | 1056 | 896 | P
indentation | 500_HK_1_1_20X.bmp | 1056 | 896 | P
indentation | 500_HK_2_1_20X.bmp | 1056 | 896 | P
indentation | 500_KH_1_1_20X.bmp | 1056 | 896 | P
indentation | 600_HBW_10_3000_1_10X.bmp | 1056 | 896 | P
indentation | 600_HBW_10_3000_1_20X.bmp | 1056 | 896 | P
indentation | 600_HBW_10_500_1_10X.bmp | 1056 | 896 | P
indentation | 600_HBW_1_30_1_10X.bmp | 1056 | 896 | P
indentation | 600_HBW_1_30_1_20X.bmp | 1056 | 896 | P
indentation | 600_HBW_1_5_1_20X.bmp | 1056 | 896 | P
indentation | 706-437_307_HMV_0.5_1_50X.bmp | 1056 | 896 | P
indentation | 706-437_307_HMV_1_1_50X.bmp | 1056 | 896 | P
indentation | 706-437_307_HMV_1_2_50X.bmp | 1056 | 896 | P
indentation | 706-437_307_HMV_2_1_50X.bmp | 1056 | 896 | P
indentation | 709-293_904_HMV_0.5_1_50X.bmp | 1056 | 896 | P
indentation | 709-293_904_HMV_1_1_50X.bmp | 1056 | 896 | P
indentation | 709-293_904_HMV_2_1_50X.bmp | 1056 | 896 | P
indentation | 710-106_708_HMV_0.5_1_50X.bmp | 1056 | 896 | P
indentation | 710-106_708_HMV_1_1_50X.bmp | 1056 | 896 | P
indentation | 710-106_708_HMV_2_1_50X.bmp | 1056 | 896 | P
indentation | 710-404_39.6_HMV_0.05_1_50X.bmp | 1056 | 896 | P
indentation | 710-655_197_HMV_0.05_1_50X.bmp | 1056 | 896 | P
indentation | 711-081_102_HMV_0.05_1_50X.bmp | 1056 | 896 | P
indentation | 711-081_102_HMV_0.5_1_50X.bmp | 1056 | 896 | P
indentation | 711-081_102_HMV_1_1_20X.bmp | 1056 | 896 | P
indentation | 711-081_102_HMV_2_1_20X.bmp | 1056 | 896 | P
indentation | 711-246_502_HMV_0.5_1_50X.bmp | 1056 | 896 | P
indentation | 711-246_502_HMV_1_1_50X.bmp | 1056 | 896 | P
indentation | 711-246_502_HMV_1_3_50X.bmp | 1056 | 896 | P
indentation | 711-246_502_HMV_2_1_50X.bmp | 1056 | 896 | P
indentation | 900_HK_0.05_1_50X.bmp | 1056 | 896 | P
indentation | 900_HK_2_1_20X.bmp | 1056 | 896 | P