## Vickers
Serial Number | Hardness scale | Position | Applied force [N] | Magnification | Manual | [7 empty] | Automatic | [7 empty] | Picture ID
[5 empty] | d1 [μm] | d2 [μm] | Average [μm] | Average [μm] | Hardness | Average hardness | Expanded standard uncertainty | Relative expanded standard uncertainty | d1 [μm] | d2 [μm] | Average [μm] | Average [μm] | Hardness | Average hardness | Expanded standard uncertainty | Relative expanded standard uncertainty
293-250 | HV 3 | 1 | 29.43 | 20X | 183.943 | 183.904 | 183.9235 | 184.5562 | 164.421 | 163.5923 | 2.5363 | 1.5504 | 184.291 | 184.428 | 184.3595 | 184.5995 | 163.709 | 163.2407 | 2.0781 | 1.273 | 293-250_160_HV_3_1_20X
293-250 | HV 3 | 2 | 29.43 | 20X | 184.644 | 185.667 | 185.1555 |  | 163.12 |  |  |  | 184.721 | 184.703 | 184.712 |  | 163.02
293-250 | HV 3 | 3 | 29.43 | 20X | 184.115 | 185.064 | 184.5895 |  | 163.236 |  |  |  | 184.334 | 185.12 | 184.727 |  | 162.993
294-367 | HV 3 | 1 | 29.43 | 20X | 135.504 | 135.211 | 135.3575 | 135.062 | 303.553 | 304.9003 | 4.2319 | 1.388 | 135.698 | 135.197 | 135.4475 | 135.0908 | 303.17 | 304.4983 | 4.3009 | 1.4124 | 294-367_300_HV_3_1_20X
294-367 | HV 3 | 2 | 29.43 | 20X | 135.058 | 134.521 | 134.7895 |  | 306.139 |  |  |  | 135.251 | 134.327 | 134.789 |  | 306.141
294-367 | HV 3 | 3 | 29.43 | 20X | 135.182 | 134.896 | 135.039 |  | 305.009 |  |  |  | 134.85 | 135.222 | 135.036 |  | 304.184
294-481 | HV 3 | 1 | 29.43 | 20X | 103.915 | 103.59 | 103.7525 | 103.873 | 516.694 | 515.499 | 7.4144 | 1.4383 | 103.96 | 104.292 | 104.126 | 104.2282 | 512.994 | 511.9953 | 7.4379 | 1.4527 | 294_481_503_HV_3_1_20X
294-481 | HV 3 | 2 | 29.43 | 20X | 103.585 | 104.006 | 103.7955 |  | 516.266 |  |  |  | 103.962 | 104.188 | 104.075 |  | 513.497
294-481 | HV 3 | 3 | 29.43 | 20X | 103.88 | 104.262 | 104.071 |  | 513.537 |  |  |  | 104.704 | 104.263 | 104.4835 |  | 509.495
216-610 | HV 3 | 1 | 29.43 | 50X | 86.998 | 87.315 | 87.1565 | 87.4893 | 732.203 | 726.6623 | 11.4188 | 1.5714 | 86.496 | 86.765 | 86.6305 | 87.2083 | 740.873 | 731.3793 | 14.3219 | 1.9582 | 216-610_701_HV_3_1_50X
216-610 | HV 3 | 2 | 29.43 | 50X | 87.835 | 87.775 | 87.805 |  | 721.427 |  |  |  | 87.981 | 88.341 | 88.161 |  | 715.613
216-610 | HV 3 | 3 | 29.43 | 50X | 87.337 | 87.676 | 87.5065 |  | 726.357 |  |  |  | 86.622 | 87.045 | 86.8335 |  | 737.652
221-101 | HV 3 | 1 | 29.43 | 50X | 78.239 | 78.245 | 78.242 | 78.1743 | 908.555 | 910.1297 | 14.1415 | 1.5538 | 78.166 | 77.845 | 78.0055 | 77.9745 | 914.084 | 914.8567 | 15.3644 | 1.6794 | 221-101_898_HV_3_1_50X
221-101 | HV 3 | 2 | 29.43 | 50X | 78.122 | 78.238 | 78.18 |  | 909.996 |  |  |  | 77.253 | 77.823 | 77.538 |  | 925.128
221-101 | HV 3 | 3 | 29.43 | 50X | 78.092 | 78.11 | 78.101 |  | 911.838 |  |  |  | 78.291 | 78.469 | 78.38 |  | 905.358
293-250 | HV 20 | 1 | 196.2 | 20X | 478.702 | 478.167 | 478.4345 | 479.2792 | 161.992 | 161.4267 | 4.9143 | 3.0443 | 478.184 | 477.252 | 477.718 | 479.3273 | 162.479 | 161.3973 | 6.1848 | 3.832
293-250 | HV 20 | 2 | 196.2 | 20X | 481.128 | 481.706 | 481.417 |  | 159.991 |  |  |  | 481.719 | 482.478 | 482.0985 |  | 159.539
293-250 | HV 20 | 3 | 196.2 | 20X | 478.297 | 477.675 | 477.986 |  | 162.297 |  |  |  | 478.515 | 477.816 | 478.1655 |  | 162.174
294-367 | HV 20 | 1 | 196.2 | 20X | 349.164 | 349.313 | 349.2385 | 349.196 | 304.015 | 304.09 | 3.7984 | 1.2491 | 349.57 | 350.015 | 349.7925 | 349.4752 | 303.054 | 303.6047 | 3.8082 | 1.2543
294-367 | HV 20 | 2 | 196.2 | 20X | 349.606 | 349.457 | 349.5315 |  | 303.506 |  |  |  | 349.852 | 349.266 | 349.559 |  | 303.458
294-367 | HV 20 | 3 | 196.2 | 20X | 348.764 | 348.872 | 348.818 |  | 304.749 |  |  |  | 348.974 | 349.174 | 349.074 |  | 304.302
294-481 | HV 20 | 1 | 196.2 | 20X | 270.186 | 269.66 | 269.923 | 270.0955 | 508.932 | 508.284 | 6.497 | 1.2782 | 270.038 | 269.777 | 269.9075 | 269.9183 | 508.993 | 508.951 | 6.313 | 1.2404 | 294_481_503_HV_20_1_20X
294-481 | HV 20 | 2 | 196.2 | 20X | 269.873 | 269.85 | 269.8615 |  | 509.164 |  |  |  | 269.547 | 269.722 | 269.6345 |  | 510.02
294-481 | HV 20 | 3 | 196.2 | 20X | 270.479 | 270.525 | 270.502 |  | 506.756 |  |  |  | 270.397 | 270.029 | 270.213 |  | 507.84
216-610 | HV 20 | 1 | 196.2 | 20X | 227.965 | 227.478 | 227.7215 | 227.4537 | 715.042 | 716.7313 | 9.5727 | 1.3356 | 227.615 | 227.823 | 227.719 | 227.1698 | 715.058 | 717.44 | 10.0422 | 1.3997 | 216-610_701_HV_20_1_20X
216-610 | HV 20 | 2 | 196.2 | 20X | 227.757 | 227.532 | 227.6445 |  | 715.526 |  |  |  | 227.073 | 226.9 | 226.9865 |  | 716.423
216-610 | HV 20 | 3 | 196.2 | 20X | 227.274 | 226.716 | 226.995 |  | 719.626 |  |  |  | 226.263 | 227.345 | 226.804 |  | 720.839
221-101 | HV 20 | 1 | 196.2 | 20X | 202.798 | 202.576 | 202.687 | 202.8217 | 902.585 | 901.3887 | 11.2763 | 1.251 | 202.346 | 203.176 | 202.761 | 202.566 | 901.926 | 903.6693 | 12.0942 | 1.3383 | 221-101_898_HV_20_1_20X
221-101 | HV 20 | 2 | 196.2 | 20X | 203.287 | 202.824 | 203.0555 |  | 899.312 |  |  |  | 203.227 | 202.373 | 202.8 |  | 901.579
221-101 | HV 20 | 3 | 196.2 | 20X | 202.939 | 202.506 | 202.7225 |  | 902.269 |  |  |  | 202.175 | 202.099 | 202.137 |  | 907.503
293-250 | HV 50 | 1 | 490.5 | 20X | 756.607 | 756.335 | 756.471 | 756.2575 | 162 | 162.1 | 6.9689 | 4.2991 | 756.809 | 757.33 | 757.0695 | 756.645 | 162 | 162.0333 | 7.3001 | 4.5053
293-250 | HV 50 | 2 | 490.5 | 20X | 758.557 | 759.2 | 758.8785 |  | 161 |  |  |  | 758.982 | 759.583 | 759.2825 |  | 160.8
293-250 | HV 50 | 3 | 490.5 | 20X | 753.114 | 753.732 | 753.423 |  | 163.3 |  |  |  | 753.544 | 753.622 | 753.583 |  | 163.3
294-367 | HV 50 | 1 | 490.5 | 20X | 550.791 | 552.234 | 551.5125 | 551.6933 | 304.8 | 304.6333 | 3.8664 | 1.2692 | 550.99 | 552.97 | 551.98 | 551.7327 | 304.3 | 304.5667 | 3.851 | 1.2644
294-367 | HV 50 | 2 | 490.5 | 20X | 552.214 | 552.175 | 552.1945 |  | 304.1 |  |  |  | 552.339 | 551.633 | 551.986 |  | 304.3
294-367 | HV 50 | 3 | 490.5 | 20X | 550.944 | 551.802 | 551.373 |  | 305 |  |  |  | 550.974 | 551.49 | 551.232 |  | 305.1
294-481 | HV 50 | 1 | 490.5 | 20X | 426.107 | 426.543 | 426.325 | 426.5372 | 510.1 | 509.6 | 5.8116 | 1.1404 | 426.539 | 426.462 | 426.5005 | 426.8152 | 509.7 | 508.9333 | 6.2624 | 1.2305
294-481 | HV 50 | 2 | 490.5 | 20X | 426.331 | 427.089 | 426.71 |  | 509.2 |  |  |  | 426.734 | 427.684 | 427.209 |  | 508
294-481 | HV 50 | 3 | 490.5 | 20X | 426.65 | 426.503 | 426.5765 |  | 509.5 |  |  |  | 427.247 | 426.225 | 426.736 |  | 509.1
216-610 | HV 50 | 1 | 490.5 | 20X | 360.468 | 360.302 | 360.385 | 360.7707 | 713.9 | 712.3667 | 9.0466 | 1.2699 | 360.391 | 360.031 | 360.211 | 360.494 | 714.6 | 713.4667 | 8.4852 | 1.1893
216-610 | HV 50 | 2 | 490.5 | 20X | 360.678 | 360.832 | 360.755 |  | 712.4 |  |  |  | 360.795 | 360.281 | 360.538 |  | 713.3
216-610 | HV 50 | 3 | 490.5 | 20X | 361.138 | 361.206 | 361.172 |  | 710.8 |  |  |  | 361.025 | 360.441 | 360.733 |  | 712.5
221-101 | HV 50 | 1 | 490.5 | 20X | 321.971 | 321.549 | 321.76 | 321.8033 | 895.5 | 895.3 | 10.2999 | 1.1504 | 321.004 | 320.734 | 320.869 | 321.2492 | 900.5 | 898.3667 | 11.1392 | 1.2399
221-101 | HV 50 | 2 | 490.5 | 20X | 321.852 | 322.044 | 321.948 |  | 894.5 |  |  |  | 321.677 | 321.204 | 321.4405 |  | 897.3
221-101 | HV 50 | 3 | 490.5 | 20X | 321.762 | 321.642 | 321.702 |  | 895.9 |  |  |  | 321.678 | 321.198 | 321.438 |  | 897.3
293-250 | HV 100 | 1 | 981 | 20X | 1066.623 | 1071.264 | 1068.9435 | 1069.2362 | 162.3 | 162.2 | 7.7835 | 4.7987 | 1066.977 | 1071.568 | 1069.2725 | 1069.524 | 162.2 | 162.1333 | 7.7406 | 4.7742
293-250 | HV 100 | 2 | 981 | 20X | 1072.357 | 1072.534 | 1072.4455 |  | 161.2 |  |  |  | 1072.775 | 1072.623 | 1072.699 |  | 161.2
293-250 | HV 100 | 3 | 981 | 20X | 1064.814 | 1067.825 | 1066.3195 |  | 163.1 |  |  |  | 1065.147 | 1068.054 | 1066.6005 |  | 163
294-367 | HV 100 | 1 | 981 | 20X | 781.723 | 781.85 | 781.7865 | 781.4813 | 303.4 | 303.6667 | 3.4297 | 1.1294 | 781.777 | 782.076 | 781.9265 | 781.7035 | 303.3 | 303.5 | 3.3931 | 1.118
294-367 | HV 100 | 2 | 981 | 20X | 780.178 | 782.457 | 781.3175 |  | 303.8 |  |  |  | 781.035 | 782.443 | 781.739 |  | 303.5
294-367 | HV 100 | 3 | 981 | 20X | 780.459 | 782.221 | 781.34 |  | 303.8 |  |  |  | 780.912 | 781.978 | 781.445 |  | 303.7
294-481 | HV 100 | 1 | 981 | 20X | 605.403 | 606.456 | 605.9295 | 605.7117 | 505.1 | 505.4667 | 7.6895 | 1.5213 | 605.16 | 606.295 | 605.7275 | 605.5448 | 505.4 | 505.7333 | 8.955 | 1.7707
294-481 | HV 100 | 2 | 981 | 20X | 605.952 | 606.62 | 606.286 |  | 504.5 |  |  |  | 606.088 | 606.642 | 606.365 |  | 504.4
294-481 | HV 100 | 3 | 981 | 20X | 604.529 | 605.31 | 604.9195 |  | 506.8 |  |  |  | 604.142 | 604.942 | 604.542 |  | 507.4
216-610 | HV 100 | 1 | 981 | 20X | 511.016 | 511.678 | 511.347 | 511.2327 | 709.2 | 709.5333 | 8.5269 | 1.2018 | 511.311 | 511.273 | 511.292 | 511.089 | 709.4 | 709.9667 | 8.7636 | 1.2344
216-610 | HV 100 | 2 | 981 | 20X | 510.923 | 510.777 | 510.85 |  | 710.6 |  |  |  | 510.818 | 510.467 | 510.6425 |  | 711.2
216-610 | HV 100 | 3 | 981 | 20X | 511.452 | 511.55 | 511.501 |  | 708.8 |  |  |  | 511.357 | 511.308 | 511.3325 |  | 709.3
221-101 | HV 100 | 1 | 981 | 20X | 455.893 | 456.194 | 456.0435 | 456.0675 | 891.7 | 891.6 | 10.0015 | 1.1217 | 455.915 | 456.254 | 456.0845 | 455.9177 | 891.5 | 892.1667 | 10.0253 | 1.1237
221-101 | HV 100 | 2 | 981 | 20X | 455.818 | 456.02 | 455.919 |  | 892.2 |  |  |  | 456.009 | 455.49 | 455.7495 |  | 892.8
221-101 | HV 100 | 3 | 981 | 20X | 456.126 | 456.354 | 456.24 |  | 890.9 |  |  |  | 455.607 | 456.231 | 455.919 |  | 892.2
Taken as the CMC of the machine
|  |  |  | HV3 - HV100 | 1+43/d

## Mirco-Vickers
Serial Number | Hardness scale | Position | Applied force [N] | Magnification | Manual | [7 empty] | Automatic | [7 empty] | Picture ID
[5 empty] | d1 [μm] | d2 [μm] | Average [μm] | Average [μm] | Hardness | Average hardness | Expanded standard uncertainty | Relative expanded standard uncertainty | d1 [μm] | d2 [μm] | Average [μm] | Average [μm] | Hardness | Average hardness | Expanded standard uncertainty | Relative expanded standard uncertainty
710-404 | HV 0,5 | 1 | 0.4905 | 50X | 44.878 | 43.552 | 44.215 | 43.3358 | 47.418 | 49.4187 | 1.0647 | 2.1544 | 44.954 | 43.802 | 44.378 | 43.4723 | 47.07 | 49.1073 | 1.0535 | 2.1453 | 710-404_39.6_HMV_0.05_1_50X
710-404 | HV 0,5 | 2 | 0.4905 | 50X | 43.978 | 43.227 | 43.6025 |  | 48.759 |  |  |  | 43.997 | 43.381 | 43.689 |  | 48.566
710-404 | HV 0,5 | 3 | 0.4905 | 50X | 42.364 | 42.016 | 42.19 |  | 52.079 |  |  |  | 42.393 | 42.307 | 42.35 |  | 51.686
711-081 | HV 0,5 | 1 | 0.4905 | 50X | 28.754 | 28.431 | 28.5925 | 28.7642 | 113.39 | 112.061 | 1.9574 | 1.7467 | 28.748 | 28.501 | 28.6245 | 28.8307 | 113.141 | 111.551 | 1.9569 | 1.7542 | 711-081_102_HMV_0.05_1_50X
711-081 | HV 0,5 | 2 | 0.4905 | 50X | 29.203 | 28.962 | 29.0825 |  | 109.601 |  |  |  | 29.39 | 28.959 | 29.1745 |  | 108.915
711-081 | HV 0,5 | 3 | 0.4905 | 50X | 29.042 | 28.193 | 28.6175 |  | 113.192 |  |  |  | 29.243 | 28.143 | 28.693 |  | 112.597
710-655 | HV 0,5 | 1 | 0.4905 | 50X | 20.888 | 20.796 | 20.842 | 20.8492 | 217.337 | 214.576 | 4.2145 | 1.9641 | 20.921 | 20.887 | 20.904 | 20.8857 | 212.139 | 212.5323 | 4.1871 | 1.9701 | 710-655_197_HMV_0.05_1_50X
710-655 | HV 0,5 | 2 | 0.4905 | 50X | 20.584 | 20.941 | 20.7625 |  | 215.041 |  |  |  | 20.441 | 21.023 | 20.732 |  | 215.674
710-655 | HV 0,5 | 3 | 0.4905 | 50X | 20.701 | 21.185 | 20.943 |  | 211.35 |  |  |  | 20.762 | 21.28 | 21.021 |  | 209.784
711-081 | HV 0,5 | 1 | 4.905 | 50X | 94.577 | 94.173 | 94.375 | 93.8983 | 104.08 | 105.1503 | 1.6828 | 1.6003 | 94.669 | 94.387 | 94.528 | 94.0422 | 103.743 | 104.827 | 1.6047 | 1.5308 | 711-081_102_HMV_0.5_1_50X
711-081 | HV 0,5 | 2 | 4.905 | 50X | 93.058 | 93.169 | 93.1135 |  | 106.919 |  |  |  | 93.436 | 93.271 | 93.3535 |  | 106.371
711-081 | HV 0,5 | 3 | 4.905 | 50X | 94.391 | 94.022 | 94.2065 |  | 104.452 |  |  |  | 94.581 | 93.909 | 94.245 |  | 104.367
706-437 | HV 0,5 | 1 | 4.905 | 50X | 54.053 | 53.972 | 54.0125 | 53.9568 | 318.225 | 318.5757 | 4.4691 | 1.4029 | 53.836 | 53.96 | 53.898 | 54.0013 | 319.106 | 317.8967 | 4.4913 | 1.4128 | 706-437_307_HMV_0.5_1_50X
706-437 | HV 0,5 | 2 | 4.905 | 50X | 53.833 | 53.645 | 53.739 |  | 320.997 |  |  |  | 53.868 | 53.824 | 53.846 |  | 319.722
706-437 | HV 0,5 | 3 | 4.905 | 50X | 54.12 | 54.118 | 54.119 |  | 316.505 |  |  |  | 53.817 | 54.703 | 54.26 |  | 314.862
711-246 | HV 0,5 | 1 | 4.905 | 50X | 42.855 | 42.925 | 42.89 | 42.603 | 503.927 | 510.7747 | 7.7559 | 1.5185 | 42.965 | 43.139 | 43.052 | 42.7943 | 500.142 | 506.258 | 8.0059 | 1.5814 | 711-246_502_HMV_0.5_1_50X
711-246 | HV 0,5 | 2 | 4.905 | 50X | 42.539 | 42.309 | 42.424 |  | 515.058 |  |  |  | 42.483 | 42.216 | 42.3495 |  | 516.86
711-246 | HV 0,5 | 3 | 4.905 | 50X | 42.506 | 42.484 | 42.495 |  | 513.339 |  |  |  | 42.723 | 43.24 | 42.9815 |  | 501.772
710-106 | HV 0,5 | 1 | 4.905 | 50X | 35.089 | 35.34 | 35.2145 | 35.1163 | 747.544 | 751.7683 | 11.9766 | 1.5931 | 35.068 | 34.808 | 34.938 | 35.0033 | 759.423 | 756.691 | 12.3385 | 1.6306 | 710-106_708_HMV_0.5_1_50X
710-106 | HV 0,5 | 2 | 4.905 | 50X | 35.316 | 35.134 | 35.225 |  | 747.098 |  |  |  | 35.396 | 35.236 | 35.316 |  | 743.253
710-106 | HV 0,5 | 3 | 4.905 | 50X | 34.912 | 34.907 | 34.9095 |  | 760.663 |  |  |  | 34.438 | 35.074 | 34.756 |  | 767.397
709-293 | HV 0,5 | 1 | 4.905 | 50X | 30.615 | 30.99 | 30.8025 | 30.7273 | 977.03 | 981.8373 | 16.2799 | 1.6581 | 30.856 | 31.007 | 30.9315 | 30.7503 | 968.929 | 980.475 | 16.5296 | 1.6859 | 709-293_904_HMV_0.5_1_50X
709-293 | HV 0,5 | 2 | 4.905 | 50X | 30.634 | 30.901 | 30.7675 |  | 979.254 |  |  |  | 30.597 | 31.047 | 30.822 |  | 975.794
709-293 | HV 0,5 | 3 | 4.905 | 50X | 30.769 | 30.455 | 30.612 |  | 989.228 |  |  |  | 30.544 | 30.451 | 30.4975 |  | 996.702
711-081 | HV 1 | 1 | 9.81 | 20X | 133.989 | 133.847 | 133.918 | 133.8632 | 103.379 | 103.4753 | 1.3348 | 1.29 | 134.227 | 134.149 | 134.188 | 133.9907 | 102.963 | 103.2693 | 1.3956 | 1.3514 | 711-081_102_HMV_1_1_20X
711-081 | HV 1 | 2 | 9.81 | 20X | 133.772 | 134.662 | 134.217 |  | 102.949 |  |  |  | 133.884 | 134.768 | 134.326 |  | 102.752
711-081 | HV 1 | 3 | 9.81 | 20X | 133.896 | 133.013 | 133.4545 |  | 104.098 |  |  |  | 133.959 | 132.957 | 133.458 |  | 104.093
706-437 | HV 1 | 1 | 9.81 | 50X | 76.986 | 78.473 | 77.7295 | 76.676 | 306.858 | 315.334 | 6.3542 | 2.0151 | 76.23 | 78.803 | 77.5165 | 76.3145 | 308.543 | 318.487 | 6.8887 | 2.1629 | 706-437_307_HMV_1_1_50X
706-437 | HV 1 | 2 | 9.81 | 50X | 76.099 | 75.238 | 75.6685 |  | 323.417 |  |  |  | 75.404 | 75.034 | 75.219 |  | 327.684 |  |  |  | 706-437_307_HMV_1_2_50X
706-437 | HV 1 | 3 | 9.81 | 50X | 76.978 | 76.282 | 76.63 |  | 315.727 |  |  |  | 76.603 | 75.813 | 76.208 |  | 319.234
711-246 | HV 1 | 1 | 9.81 | 50X | 61.233 | 60.571 | 60.902 | 60.518 | 499.858 | 506.2687 | 7.449 | 1.4713 | 61.339 | 60.506 | 60.9225 | 60.277 | 499.513 | 510.365 | 8.0699 | 1.5812 | 711-246_502_HMV_1_1_50X
711-246 | HV 1 | 2 | 9.81 | 50X | 60.38 | 60.759 | 60.5695 |  | 505.361 |  |  |  | 59.748 | 60.131 | 59.9395 |  | 516.049
711-246 | HV 1 | 3 | 9.81 | 50X | 60.709 | 59.456 | 60.0825 |  | 513.587 |  |  |  | 60.284 | 59.654 | 59.969 |  | 515.533 |  |  |  | 711-246_502_HMV_1_3_50X
710-106 | HV 1 | 1 | 9.81 | 50X | 50.597 | 50.739 | 50.668 | 50.3497 | 722.175 | 731.3917 | 10.7947 | 1.4759 | 50.109 | 50.502 | 50.3055 | 50.0255 | 732.635 | 740.9113 | 10.9011 | 1.4713 | 710-106_708_HMV_1_1_50X
710-106 | HV 1 | 2 | 9.81 | 50X | 50.4 | 50.273 | 50.3365 |  | 731.718 |  |  |  | 50.242 | 49.873 | 50.0575 |  | 739.912
710-106 | HV 1 | 3 | 9.81 | 50X | 49.855 | 50.234 | 50.0445 |  | 740.282 |  |  |  | 49.512 | 49.915 | 49.7135 |  | 750.187
709-293 | HV 1 | 1 | 9.81 | 50X | 45.061 | 45.141 | 45.101 | 45.1108 | 911.46 | 911.0723 | 13.23 | 1.4521 | 43.991 | 44.591 | 44.291 | 44.3858 | 945.102 | 941.148 | 14.2215 | 1.5111 | 709-293_904_HMV_1_1_50X
709-293 | HV 1 | 2 | 9.81 | 50X | 45.165 | 44.858 | 45.0115 |  | 915.088 |  |  |  | 44.133 | 44.203 | 44.168 |  | 950.373
709-293 | HV 1 | 3 | 9.81 | 50X | 45.172 | 45.268 | 45.22 |  | 906.669 |  |  |  | 44.82 | 44.577 | 44.6985 |  | 927.969
711-081 | HV 2 | 1 | 19.62 | 20X | 190.789 | 189.096 | 189.9425 | 189.3157 | 102.777 | 103.472 | 2.637 | 2.5485 | 191.047 | 189.242 | 190.1445 | 189.6583 | 102.558 | 103.097 | 2.556 | 2.4793 | 711-081_102_HMV_2_1_20X
711-081 | HV 2 | 2 | 19.62 | 20X | 188.815 | 186.379 | 187.597 |  | 105.363 |  |  |  | 188.96 | 187.083 | 188.0215 |  | 104.887
711-081 | HV 2 | 3 | 19.62 | 20X | 190.903 | 189.912 | 190.4075 |  | 102.276 |  |  |  | 191.346 | 190.272 | 190.809 |  | 101.846
706-437 | HV 2 | 1 | 19.62 | 20X | 108.2 | 108.943 | 108.5715 | 108.0538 | 314.563 | 317.597 | 4.4175 | 1.3909 | 108.206 | 109.032 | 108.619 | 108.0802 | 314.288 | 317.443 | 4.4086 | 1.3888 | 706-437_307_HMV_2_1_50X
706-437 | HV 2 | 2 | 19.62 | 20X | 107.881 | 108.039 | 107.96 |  | 318.137 |  |  |  | 107.664 | 108.136 | 107.9 |  | 318.491
706-437 | HV 2 | 3 | 19.62 | 20X | 107.52 | 107.74 | 107.63 |  | 320.091 |  |  |  | 107.642 | 107.801 | 107.7215 |  | 319.55
711-246 | HV 2 | 1 | 19.62 | 50X | 84.549 | 84.795 | 84.672 | 84.9267 | 517.202 | 514.112 | 6.5828 | 1.2804 | 84.358 | 84.768 | 84.563 | 84.8642 | 518.536 | 514.842 | 6.6849 | 1.2984 | 711-246_502_HMV_2_1_50X
711-246 | HV 2 | 2 | 19.62 | 50X | 85.104 | 85.004 | 85.054 |  | 512.567 |  |  |  | 85.039 | 84.949 | 84.994 |  | 513.194
711-246 | HV 2 | 3 | 19.62 | 50X | 85.007 | 85.101 | 85.054 |  | 512.567 |  |  |  | 84.98 | 85.091 | 85.0355 |  | 512.796
710-106 | HV 2 | 1 | 19.62 | 50X | 70.591 | 71.012 | 70.8015 | 70.8057 | 739.699 | 739.6133 | 9.5237 | 1.2877 | 70.467 | 71.053 | 70.76 | 70.5585 | 740.567 | 744.817 | 9.8039 | 1.3163 | 710-106_708_HMV_2_1_50X
710-106 | HV 2 | 2 | 19.62 | 50X | 70.983 | 70.784 | 70.8835 |  | 737.988 |  |  |  | 71.025 | 70.049 | 70.537 |  | 745.256
710-106 | HV 2 | 3 | 19.62 | 50X | 70.853 | 70.611 | 70.732 |  | 741.153 |  |  |  | 70.55 | 70.207 | 70.3785 |  | 748.628
709-293 | HV 2 | 1 | 19.62 | 50X | 63.552 | 63.767 | 63.6595 | 63.2582 | 914.984 | 926.7043 | 13.4406 | 1.4504 | 63.73 | 63.764 | 63.747 | 63.1747 | 912.473 | 929.244 | 14.9652 | 1.6105 | 709-293_904_HMV_2_1_50X
709-293 | HV 2 | 2 | 19.62 | 50X | 63.336 | 63.174 | 63.255 |  | 926.723 |  |  |  | 63.35 | 63.133 | 63.2415 |  | 927.104
709-293 | HV 2 | 3 | 19.62 | 50X | 62.592 | 63.128 | 62.86 |  | 938.406 |  |  |  | 62.142 | 62.929 | 62.5355 |  | 948.155
Taken as the CMC of the machine
|  |  |  | HV0,05 - HV2 | 1+20/d

## Brinell
Serial Number | Hardness scale | Position | Ball diameter [mm] | Applied force [N] | Magnification | Manual | [7 empty] | Automatic | [7 empty] | Picture ID
[6 empty] | d1 [μm] | d2 [μm] | Average [μm] | Average [μm] | Hardness | Average hardness | Expanded standard uncertainty | Relative expanded standard uncertainty | d1 [μm] | d2 [μm] | Average [μm] | Average [μm] | Hardness | Average hardness | Expanded standard uncertainty | Relative expanded standard uncertainty
293-250 | HBW 1/5 | 1 | 1 | 49.05 | 20x | 209.976 | 211.922 | 210.949 | 210.4625 | 141.497 | 142.1683 | 3.9649 | 2.7889 | 211.318 | 213.35 | 212.334 | 211.8828 | 139.636 | 140.241 | 3.3499 | 2.3886 | 100_HBW_1_5_1_20X
293-250 | HBW 1/5 | 2 | 1 | 49.05 | 20x | 211.436 | 211.192 | 211.314 |  | 141.003 |  |  |  | 212.215 | 212.275 | 212.245 |  | 139.754
293-250 | HBW 1/5 | 3 | 1 | 49.05 | 20x | 209.49 | 208.759 | 209.1245 |  | 144.005 |  |  |  | 211.252 | 210.887 | 211.0695 |  | 141.333
294-367 | HBW 1/5 | 1 | 1 | 49.05 | 20x | 154.258 | 154.988 | 154.623 | 154.5825 | 264.758 | 264.8983 | 6.8031 | 2.5682 | 155.476 | 155.559 | 155.5175 | 155.1857 | 261.701 | 262.832 | 6.8719 | 2.6145 | 300_HBW_1_5_1_20X
294-367 | HBW 1/5 | 2 | 1 | 49.05 | 20x | 154.501 | 154.988 | 154.7445 |  | 264.34 |  |  |  | 155.086 | 155.477 | 155.2815 |  | 262.502
294-367 | HBW 1/5 | 3 | 1 | 49.05 | 20x | 154.745 | 154.015 | 154.38 |  | 265.597 |  |  |  | 155.118 | 154.398 | 154.758 |  | 264.293
290-736 | HBW 1/5 | 1 | 1 | 49.05 | 20x | 115.815 | 116.059 | 115.937 | 116.9922 | 472.178 | 463.921 | 19.6775 | 4.2416 | 117.761 | 117.033 | 117.397 | 117.9267 | 460.467 | 456.5573 | 19.075 | 4.178 | 600_HBW_1_5_1_20X
290-736 | HBW 1/5 | 2 | 1 | 49.05 | 20x | 119.224 | 119.222 | 119.223 |  | 446.421 |  |  |  | 120.442 | 119.595 | 120.0185 |  | 440.498
290-736 | HBW 1/5 | 3 | 1 | 49.05 | 20x | 115.086 | 116.547 | 115.8165 |  | 473.164 |  |  |  | 115.698 | 117.031 | 116.3645 |  | 468.707
293-250 | HBW 1/30 | 1 | 1 | 294.3 | 20x | 466.565 | 463.586 | 465.0755 | 464.74 | 166.519 | 166.7753 | 2.7303 | 1.6371 | 470.313 | 465.067 | 467.69 | 466.62 | 164.542 | 165.3487 | 3.4346 | 2.0772
293-250 | HBW 1/30 | 2 | 1 | 294.3 | 20x | 464.745 | 464.963 | 464.854 |  | 166.688 |  |  |  | 465.691 | 466.693 | 466.192 |  | 165.671
293-250 | HBW 1/30 | 3 | 1 | 294.3 | 20x | 463.435 | 465.146 | 464.2905 |  | 167.119 |  |  |  | 465.563 | 466.393 | 465.978 |  | 165.833
294-367 | HBW 1/30 | 1 | 1 | 294.3 | 20x | 351.556 | 352.297 | 351.9265 | 350.6645 | 298.638 | 300.8717 | 7.3022 | 2.427 | 352.049 | 353.004 | 352.5265 | 351.1102 | 297.587 | 300.084 | 7.5794 | 2.5258
294-367 | HBW 1/30 | 2 | 1 | 294.3 | 20x | 350.291 | 350.468 | 350.3795 |  | 301.371 |  |  |  | 350.619 | 350.478 | 350.5485 |  | 301.072
294-367 | HBW 1/30 | 3 | 1 | 294.3 | 20x | 349.477 | 349.898 | 349.6875 |  | 302.606 |  |  |  | 349.884 | 350.627 | 350.2555 |  | 301.593
290-736 | HBW 1/30 | 1 | 1 | 294.3 | 20x | 254.522 | 254.952 | 254.737 | 254.9415 | 579.109 | 578.167 | 11.6423 | 2.0137 | 257.941 | 255.113 | 256.527 | 255.4053 | 570.918 | 576.049 | 14.0586 | 2.4405
290-736 | HBW 1/30 | 2 | 1 | 294.3 | 20x | 254.446 | 255.036 | 254.741 |  | 579.091 |  |  |  | 254.616 | 255.125 | 254.8705 |  | 578.494
290-736 | HBW 1/30 | 3 | 1 | 294.3 | 20x | 255.452 | 255.241 | 255.3465 |  | 576.301 |  |  |  | 254.809 | 254.828 | 254.8185 |  | 578.735
812-006 | HBW 10/500 | 1 | 10 | 4905 | 20x | 2201.248 | 2201.576 | 2201.412 | 2203.1337 | 129.794 | 129.5897 | 8.9941 | 6.9405 | 2231.69 | 2244.673 | 2238.1815 | 2234.308 | 125.511 | 125.954 | 10.7219 | 8.5126
812-006 | HBW 10/500 | 2 | 10 | 4905 | 20x | 2208.315 | 2208.155 | 2208.235 |  | 128.983 |  |  |  | 2239.174 | 2216.757 | 2227.9655 |  | 126.68
812-006 | HBW 10/500 | 3 | 10 | 4905 | 20x | 2199.922 | 2199.586 | 2199.754 |  | 129.992 |  |  |  | 2231.057 | 2242.497 | 2236.777 |  | 125.671
810-202 | HBW 10/500 | 1 | 10 | 4905 | 20x | 1687.008 | 1687.041 | 1687.0245 | 1686.2058 | 222.152 | 222.3693 | 4.7157 | 2.1207 | 1710.994 | 1714.631 | 1712.8125 | 1710.5187 | 215.464 | 216.0637 | 34.9003 | 16.1528
810-202 | HBW 10/500 | 2 | 10 | 4905 | 20x | 1684.493 | 1685.233 | 1684.863 |  | 222.726 |  |  |  | 1677.755 | 1720.223 | 1698.989 |  | 219.011
810-202 | HBW 10/500 | 3 | 10 | 4905 | 20x | 1687.056 | 1686.404 | 1686.73 |  | 222.23 |  |  |  | 1729.945 | 1709.564 | 1719.7545 |  | 213.716
8301-50DW | HBW 10/500 | 1 | 10 | 4905 | 10x | 1206.303 | 1204.281 | 1205.292 | 1204.3738 | 436.763 | 437.4327 | 10.3037 | 2.3555 | 1265.915 | 1195.979 | 1230.947 | 1239.9103 | 418.681 | 412.87 | 134.7277 | 32.632
8301-50DW | HBW 10/500 | 2 | 10 | 4905 | 10x | 1203.107 | 1202.59 | 1202.8485 |  | 438.546 |  |  |  | 1251.314 | 1277.344 | 1264.329 |  | 396.78
8301-50DW | HBW 10/500 | 3 | 10 | 4905 | 10x | 1205.397 | 1204.565 | 1204.981 |  | 436.989 |  |  |  | 1237.417 | 1211.493 | 1224.455 |  | 423.149
812-006 | HBW 10/3000 | 1 | 10 | 29430 | 20x | 4832.528 | 4831.871 | 4832.1995 | 4840.7873 | 153.449 | 152.8677 | 21.8225 | 14.2754 | 4861.851 | 4842.237 | 4852.044 | 4861.091 | 152.108 | 151.5053 | 38.8338 | 25.632
812-006 | HBW 10/3000 | 2 | 10 | 29430 | 20x | 4847.967 | 4853.497 | 4850.732 |  | 152.196 |  |  |  | 4868.731 | 4892.226 | 4880.4785 |  | 150.213
812-006 | HBW 10/3000 | 3 | 10 | 29430 | 20x | 4834.603 | 4844.258 | 4839.4305 |  | 152.958 |  |  |  | 4845.574 | 4855.927 | 4850.7505 |  | 152.195
810-202 | HBW 10/3000 | 1 | 10 | 29430 | 20x | 3807.15 | 3805.587 | 3806.3685 | 3804.9168 | 253.796 | 253.998 | 11.9262 | 4.6954 | 3809.201 | 3832.075 | 3820.638 | 3823.7332 | 251.827 | 251.411 | 60.0155 | 23.8715
810-202 | HBW 10/3000 | 2 | 10 | 29430 | 20x | 3800.624 | 3802.309 | 3801.4665 |  | 254.478 |  |  |  | 3838.127 | 3843.264 | 3840.6955 |  | 249.095
810-202 | HBW 10/3000 | 3 | 10 | 29430 | 20x | 3807.1 | 3806.731 | 3806.9155 |  | 253.72 |  |  |  | 3809.927 | 3809.805 | 3809.866 |  | 253.311
8301-50DW | HBW 10/3000 | 1 | 10 | 29430 | 20x | 2655.324 | 2657.826 | 2656.575 | 2649.1663 | 531.68 | 534.7197 | 52.634 | 9.8433 | 2657.666 | 2668.334 | 2663 | 2655.9918 | 529.069 | 531.926 | 61.5446 | 11.5701
8301-50DW | HBW 10/3000 | 2 | 10 | 29430 | 20x | 2642.302 | 2648.291 | 2645.2965 |  | 536.308 |  |  |  | 2663.775 | 2650.206 | 2656.9905 |  | 531.51
8301-50DW | HBW 10/3000 | 3 | 10 | 29430 | 20x | 2642.272 | 2648.983 | 2645.6275 |  | 536.171 |  |  |  | 2643.732 | 2652.238 | 2647.985 |  | 535.199
Taken as the CMC of the machine
|  | HBW 1/5 - HBW 1/30 |  |  | 1+240/d
|  | HBW 10/500 - HBW 10/3000 |  |  | 1+1200/d

## Knoop
Serial Number | Hardness (on block) | Position | Applied force [N] | Magnification | Manual |  | Automatic |  | Δ [μm] | Δ [%]
[5 empty] | d [μm] | Hardness | d [μm] | Hardness
711-081 | HK 0.05 | 1 | 0.4905 | 50x | 79.238
711-081 | HK 0.05 | 2 | 0.4905 | 50x | 81.54
711-081 | HK 0.05 | 3 | 0.4905 | 50x | 81.6
711-246 | HK 0.05 | 1 | 0.4905 | 50x | 36.686
711-246 | HK 0.05 | 2 | 0.4905 | 50x | 36.54
711-246 | HK 0.05 | 3 | 0.4905 | 50x | 36.103
709-293 | HK 0.05 | 1 | 0.4905 | 50x | 28.275
709-293 | HK 0.05 | 2 | 0.4905 | 50x | 28.515
709-293 | HK 0.05 | 3 | 0.4905 | 50x | 28.493
711-081 | HK 1 | 1 | 9.81 | 10x | 372.77
711-081 | HK 1 | 2 | 9.81 | 10x | 372.321
711-081 | HK 1 | 3 | 9.81 | 10x | 370.76
711-081 | HK 1 | 1 | 9.81 | 20x | 372.629
711-081 | HK 1 | 2 | 9.81 | 20x | 372.141
711-081 | HK 1 | 3 | 9.81 | 20x | 370.888
711-246 | HK 1 | 1 | 9.81 | 20x | 168.648
711-246 | HK 1 | 2 | 9.81 | 20x | 168.246
711-246 | HK 1 | 3 | 9.81 | 20x | 168.975
709-293 | HK 1 | 1 | 9.81 | 20x | 130.027
709-293 | HK 1 | 2 | 9.81 | 20x | 130.271
709-293 | HK 1 | 3 | 9.81 | 20x | 130.674
711-081 | HK 2 | 1 | 19.62 | 10x | 530.591
711-081 | HK 2 | 2 | 19.62 | 10x | 533.922
711-081 | HK 2 | 3 | 19.62 | 10x | 531.51
711-081 | HK 2 | 1 | 19.62 | 20x | 529.435
711-081 | HK 2 | 2 | 19.62 | 20x | 532.945
711-081 | HK 2 | 3 | 19.62 | 20x | 530.498
711-246 | HK 2 | 1 | 19.62 | 20x | 238.966
711-246 | HK 2 | 2 | 19.62 | 20x | 238.601
711-246 | HK 2 | 3 | 19.62 | 20x | 238.283
709-293 | HK 2 | 1 | 19.62 | 20x | 185.244
709-293 | HK 2 | 2 | 19.62 | 20x | 185.156
709-293 | HK 2 | 3 | 19.62 | 20x | 186.143